## Scale Calibration RMs
[6 empty] | avg | [6 empty] | avg | 9.2752 | 40.504 |  |  | 4.3679 | 5.0953
[6 empty] | stdev | [6 empty] | stdev | 0.19434 | 1.0666 |  |  | 0.11912 | 0.13895
[6 empty] | RSD (%) | [6 empty] | RSD (%) | 2.0953 | 2.6332 |  |  | 2.7271 | 2.7271
[6 empty] | known | [6 empty] | known | 9.52 | 40.81 |  |  | 4.2868 | 5.0021
[6 empty] | accuracy | [6 empty] | accuracy | -0.24484 | -0.306 |  |  | 0.081172 | 0.093261
[6 empty] | n | [6 empty] | n | 24 | 24 |  |  | 24 | 24
[6 empty] | SRM | USGS 41a |  | USGS 41a |  |  |  | SRM | USGS 41a | USGS 41a |  |  | USGS 41a | USGS 41a
EA-irms proc comp template
[6 empty] | RM | d15N | ± | d13C | ± |  |  |  | N% | C% |  |  | C:N (mass) | C:N (molar) | Material | Updated
[6 empty] | USGS 40 | -4.52 | 0.12 | -26.39 | 0.09 |  |  |  | 9.52 | 40.81 |  |  | 4.2868 | 5.0021 | L-glutamic acid | 2014-10-29
[6 empty] | USGS 41a | 47.55 | 0.15 | 36.55 | 0.08 |  |  |  | 9.52 | 40.81 |  |  | 4.2868 | 5.0021 | L-glutamic acid | 2023-06-16
[6 empty] | NIST 1570a | -0.46 | 0.3 | -27.17 | 0.19 |  |  |  | 5.57 | 38.21 |  |  | 6.86 | 8.0046 | spinach leaves | 2014-10-29
[6 empty] | NIST 1577b | 7.83 | 0.16 | -21.69 | 0.14 |  |  |  | 9.95 | 48.04 |  |  | 4.8281 | 5.6338 | bovine liver | 2014-10-29
[6 empty] | EM B2151 | 4.42 | 0.29 | -26.27 | 0.15 |  |  |  | 0.62 | 9.15 |  |  | 14.7581 | 17.2206 | sediment | 2014-10-29
[6 empty] | EM B2153 | 7.3 | 0.13 | -26.66 | 0.24 |  |  |  | 0.133 | 1.61 |  |  | 12.1053 | 14.1252 | soil | 2014-10-29
[6 empty] | NIST 2702 | pending | pending | pending | pending |  |  |  | 0.27 | 2.9 |  |  | 10.7407 | 12.533 | sediment (marine) | 2016-06-22
[6 empty] | LRM JGC20C | 2.268 | 0.154 | -24.492 | 0.085 |  |  |  | 0.13 | 1.16 |  |  | 8.89 | 10.3734 | sediment (marine) | 2016-06-21
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | d15N NIST 8573/8574 corrected (AT-Air) | Ampl 44 | d13C NIST 8573/8574 corrected (VPDB) | IRMS comment | mass (mg) | N (mg) | C (mg) | N% | C% | N (µmol) | C (µmol) | C:N (mass) | C:N (molar) | EA comment
Acquisition-0114.cf | 2023/09/18 09:07:50 | USGS 40 | USGS 40 | 0 | 0 | 739 | -4.4824 | 1568 | -26.3971 | QC OK | 0.165 | 0.015004 | 0.063908 | 9.0932 | 38.7319 | 1.0709 | 5.3212 | 4.2594 | 4.9687 | QC OK |  | N | N
Acquisition-0115.cf | 2023/09/18 09:14:57 | USGS 40 | USGS 40 | 0 | 0 | 3902 | -4.5575 | 8709 | -26.3828 | QC OK | 0.89 | 0.083641 | 0.35774 | 9.3979 | 40.1952 | 5.9701 | 29.7866 | 4.2771 | 4.9893 | QC OK |  | N | N
Acquisition-0116.cf | 2023/09/18 09:22:03 | USGS 41a | USGS 41a | 0 | 0 | 837 | 47.5764 | 1811 | 36.4984 | QC OK | 0.189 | 0.017104 | 0.073888 | 9.0497 | 39.0942 | 1.2208 | 6.1522 | 4.32 | 5.0394 | QC OK |  | N | N
Acquisition-0117.cf | 2023/09/18 09:29:10 | USGS 41a | USGS 41a | 0 | 0 | 4018 | 47.5235 | 9011 | 36.6015 | QC OK | 0.908 | 0.085927 | 0.36924 | 9.4634 | 40.6657 | 6.1333 | 30.7448 | 4.2972 | 5.0128 | QC OK |  | N | N
Acquisition-0141.cf | 2023/09/18 12:19:51 | USGS 40 | USGS 40 | 0 | 0 | 2895 | -4.4914 | 6662 | -26.3303 | QC OK | 0.661 | 0.061799 | 0.27292 | 9.3493 | 41.2882 | 4.4111 | 22.724 | 4.4162 | 5.1516 | QC OK |  | N | N
Acquisition-0142.cf | 2023/09/18 12:26:57 | USGS 41a | USGS 41a | 0 | 0 | 3213 | 47.5266 | 7315 | 36.558 | QC OK | 0.727 | 0.068681 | 0.30002 | 9.4472 | 41.2679 | 4.9023 | 24.9806 | 4.3682 | 5.0957 | QC OK |  | N | N
Acquisition-0146.cf | 2023/09/27 09:51:39 | USGS 40 | USGS 40 | 0 | 0 | 982 | -4.5745 | 2154 | -26.3754 | QC OK | 0.221 | 0.019758 | 0.088267 | 8.9402 | 39.9396 | 1.4103 | 7.3494 | 4.4674 | 5.2114 | QC OK |  | N | N
Acquisition-0147.cf | 2023/09/27 09:58:47 | USGS 40 | USGS 40 | 0 | 0 | 3750 | -4.463 | 8188 | -26.4039 | QC OK | 0.844 | 0.078958 | 0.33737 | 9.3552 | 39.9732 | 5.6359 | 28.0911 | 4.2728 | 4.9844 | QC OK |  | N | N
Acquisition-0148.cf | 2023/09/27 10:05:54 | USGS 41a | USGS 41a | 0 | 0 | 3799 | 47.3014 | 8509 | 36.6963 | QC OK | 0.85 | 0.07989 | 0.3503 | 9.3988 | 41.2123 | 5.7024 | 29.1677 | 4.3848 | 5.115 | QC OK |  | N | N
Acquisition-0149.cf | 2023/09/27 10:13:00 | USGS 41a | USGS 41a | 0 | 0 | 694 | 47.7961 | 1493 | 36.403 | QC OK | 0.153 | 0.013416 | 0.060787 | 8.7687 | 39.7303 | 0.95761 | 5.0614 | 4.5309 | 5.2854 | QC OK |  | N | N
Acquisition-0173.cf | 2023/09/27 13:03:55 | USGS 40 | USGS 40 | 0 | 0 | 1958 | -4.6051 | 4600 | -26.2349 | QC OK | 0.447 | 0.040848 | 0.1912 | 9.1383 | 42.7744 | 2.9157 | 15.9202 | 4.6808 | 5.4603 | QC OK |  | N | N
Acquisition-0174.cf | 2023/09/27 13:11:01 | USGS 41a | USGS 41a | 0 | 0 | 1781 | 47.2528 | 4172 | 34.9953 | QC OK | 0.404 | 0.036908 | 0.17177 | 9.1356 | 42.5168 | 2.6344 | 14.3021 | 4.6539 | 5.429 | QC OK |  | N | N
Acquisition-0178.cf | 2023/09/28 09:48:46 | USGS 40 | USGS 40 | 0 | 0 | 4117 | -4.5925 | 9079 | -26.3803 | QC OK | 0.926 | 0.087266 | 0.37212 | 9.424 | 40.1859 | 6.2289 | 30.9843 | 4.2642 | 4.9743 | QC OK |  | N | N
Acquisition-0179.cf | 2023/09/28 09:55:52 | USGS 40 | USGS 40 | 0 | 0 | 818 | -4.4464 | 1766 | -26.3995 | QC OK | 0.181 | 0.016763 | 0.071439 | 9.2615 | 39.4688 | 1.1965 | 5.9483 | 4.2616 | 4.9713 | QC OK |  | N | N
Acquisition-0180.cf | 2023/09/28 10:02:59 | USGS 41a | USGS 41a | 0 | 0 | 804 | 47.3944 | 1729 | 36.4603 | QC OK | 0.177 | 0.016437 | 0.069759 | 9.2862 | 39.4117 | 1.1732 | 5.8084 | 4.2441 | 4.9509 | QC OK |  | N | N
Acquisition-0181.cf | 2023/09/28 10:10:06 | USGS 41a | USGS 41a | 0 | 0 | 4056 | 47.7045 | 9037 | 36.6395 | QC OK | 0.899 | 0.085626 | 0.36801 | 9.5246 | 40.9358 | 6.1118 | 30.6422 | 4.2979 | 5.0136 | QC OK |  | N | N
Acquisition-0205.cf | 2023/09/28 13:00:53 | USGS 40 | USGS 40 | 0 | 0 | 1261 | -4.5882 | 2874 | -26.4835 | QC OK | 0.285 | 0.026395 | 0.11689 | 9.2614 | 41.0125 | 1.884 | 9.7324 | 4.4283 | 5.1657 | QC OK |  | N | N
Acquisition-0206.cf | 2023/09/28 13:08:01 | USGS 41a | USGS 41a | 0 | 0 | 1633 | 47.4628 | 3724 | 36.2354 | QC OK | 0.368 | 0.034446 | 0.15153 | 9.3604 | 41.1772 | 2.4587 | 12.6172 | 4.3991 | 5.1316 | QC OK |  | N | N
Acquisition-0109.cf | 2023/10/03 09:46:28 | USGS 40 | USGS 40 | 0 | 0 | 789 | -4.3489 | 1759 | -26.3949 | QC OK | 0.175 | 0.015842 | 0.067959 | 9.0527 | 38.8339 | 1.1308 | 5.6586 | 4.2898 | 5.0041 | QC OK |  | N | N
Acquisition-0110.cf | 2023/10/03 09:53:35 | USGS 40 | USGS 40 | 0 | 0 | 4266 | -4.6897 | 9880 | -26.3848 | QC OK | 0.956 | 0.089697 | 0.38493 | 9.3825 | 40.2645 | 6.4024 | 32.0507 | 4.2914 | 5.0061 | QC OK |  | N | N
Acquisition-0111.cf | 2023/10/03 10:00:42 | USGS 41a | USGS 41a | 0 | 0 | 808 | 47.6369 | 1813 | 36.4602 | QC OK | 0.176 | 0.016129 | 0.069811 | 9.1645 | 39.6654 | 1.1513 | 5.8127 | 4.3282 | 5.0489 | QC OK |  | N | N
Acquisition-0112.cf | 2023/10/03 10:07:49 | USGS 41a | USGS 41a | 0 | 0 | 4317 | 47.4616 | 9990 | 36.6395 | QC OK | 0.952 | 0.090507 | 0.38828 | 9.507 | 40.7854 | 6.4602 | 32.3295 | 4.29 | 5.0044 | QC OK |  | N | N
Acquisition-0136.cf | 2023/10/03 12:58:41 | USGS 40 | USGS 40 | 0 | 0 | 2748 | -4.5264 | 6587 | -25.8689 | QC OK | 0.624 | 0.058367 | 0.25893 | 9.3537 | 41.4953 | 4.1661 | 21.5596 | 4.4362 | 5.175 | QC OK |  | N | N
Acquisition-0137.cf | 2023/10/03 13:05:47 | USGS 41a | USGS 41a | 0 | 0 | 3288 | 47.225 | 7771 | 36.8861 | QC OK | 0.734 | 0.069639 | 0.30439 | 9.4876 | 41.4699 | 4.9707 | 25.3446 | 4.3709 | 5.0988 | QC OK |  | N | N

## QC RM
[6 empty] | avg | 6.0385 |  | -24.5758 |  |  |  |  | 0.25278 | 3.0972 |  |  | 12.2562 | 14.2972
[6 empty] | stdev | 0.10183 |  | 0.20325 |  |  |  |  | 0.0054961 | 0.08765 |  |  | 0.38349 | 0.44735
[6 empty] | RSD | [7 empty] | 2.174% | 2.83% |  |  | 3.129% | 3.129%
[6 empty] | known | 5.95 |  | -24.69 |  |  |  |  | 0.26 | 3.11 |  |  | 11.86 | 13.84
[6 empty] | accuracy | 0.088467 |  | 0.11423 |  |  |  |  | -0.007218 | -0.012833 |  |  | 0.39619 | 0.45718
[6 empty] | n | 20 |  | 20 |  |  |  |  | 20 | 20 |  |  | 20 | 20
[6 empty] | SRM | NIST 2702 |  | NIST 2702 |  |  |  |  | NIST 2702 | NIST 2702 |  |  | NIST 2702 | NIST 2702
EA-irms proc comp template
[6 empty] | RM | d15N | ± | d13C | ± |  |  |  | N% | C% |  |  | C:N (mass) | C:N (molar) | Material | Updated
[6 empty] | USGS 40 | -4.52 | 0.12 | -26.39 | 0.09 |  |  |  | 9.52 | 40.81 |  |  | 4.2868 | 5.0021 | L-glutamic acid | 2014-10-29
[6 empty] | USGS 41a | 47.55 | 0.15 | 36.55 | 0.08 |  |  |  | 9.52 | 40.81 |  |  | 4.2868 | 5.0021 | L-glutamic acid | 2023-06-16
[6 empty] | NIST 1570a | -0.46 | 0.3 | -27.17 | 0.19 |  |  |  | 5.57 | 38.21 |  |  | 6.86 | 8.0046 | spinach leaves | 2014-10-29
[6 empty] | NIST 1577b | 7.83 | 0.16 | -21.69 | 0.14 |  |  |  | 9.95 | 48.04 |  |  | 4.8281 | 5.6338 | bovine liver | 2014-10-29
[6 empty] | EM B2151 | 4.42 | 0.29 | -26.27 | 0.15 |  |  |  | 0.62 | 9.15 |  |  | 14.7581 | 17.2206 | sediment | 2014-10-29
[6 empty] | EM B2153 | 7.3 | 0.13 | -26.66 | 0.24 |  |  |  | 0.133 | 1.61 |  |  | 12.1053 | 14.1252 | soil | 2014-10-29
[6 empty] | NIST 2702 | 5.95 | 0.25 | -24.69 | 0.06 |  |  |  | 0.26 | 3.11 |  |  | 11.86 | 13.84 | sediment (marine) | 2022-06-13
[6 empty] | LRM JGC20C | 2.268 | 0.154 | -24.492 | 0.085 |  |  |  | 0.13 | 1.16 |  |  | 8.89 | 10.3734 | sediment (marine) | 2016-06-21
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | d15N NIST 8573/8574 corrected (AT-Air) | Ampl 44 | d13C NIST 8573/8574 corrected (VPDB) | IRMS comment | mass (mg) | N (mg) | C (mg) | N% | C% | N (µmol) | C (µmol) | C:N (mass) | C:N (molar) | EA comment
Acquisition-0118.cf | 2023/09/18 09:36:16 | NIST 2702 | NIST 2702 | 0 | 0 | 548 | 6.2919 | 3075 | -24.6064 | QC OK | 4.238 | 0.010589 | 0.12731 | 0.24985 | 3.0039 | 0.75581 | 10.5999 | 12.0226 | 14.0247 | QC OK |  | N | N
Acquisition-0119.cf | 2023/09/18 09:43:23 | NIST 2702 | NIST 2702 | 0 | 0 | 928 | 6.0654 | 5244 | -24.6601 | QC OK | 7.487 | 0.018905 | 0.22243 | 0.2525 | 2.9709 | 1.3494 | 18.5206 | 11.7659 | 13.7252 | QC OK |  | N | N
Acquisition-0138.cf | 2023/09/18 11:58:32 | NIST 2702 | NIST 2702 | 0 | 0 | 707 | 6.0652 | 4277 | -24.6959 | QC OK | 5.46 | 0.013945 | 0.17524 | 0.25541 | 3.2096 | 0.99538 | 14.5914 | 12.5665 | 14.6591 | QC OK |  | N | N
Acquisition-0139.cf | 2023/09/18 12:05:38 | NIST 2702 | NIST 2702 | 0 | 0 | 1227 | 6.0247 | 7433 | -24.6886 | QC OK | 9.838 | 0.02518 | 0.30653 | 0.25595 | 3.1158 | 1.7973 | 25.5232 | 12.1737 | 14.201 | QC OK |  | N | N
Acquisition-0140.cf | 2023/09/18 12:12:44 | NIST 2702 | NIST 2702 | 0 | 0 | 949 | 6.1039 | 5769 | -24.5965 | QC OK | 7.535 | 0.019182 | 0.23703 | 0.25458 | 3.1457 | 1.3692 | 19.7359 | 12.3566 | 14.4143 | QC OK |  | N | N
Acquisition-0150.cf | 2023/09/27 10:20:09 | NIST 2702 | NIST 2702 | 0 | 0 | 526 | 6.0101 | 2963 | -24.4605 | QC OK | 3.991 | 0.0096145 | 0.12255 | 0.24091 | 3.0706 | 0.68626 | 10.2039 | 12.7461 | 14.8687 | QC OK |  | N | N
Acquisition-0151.cf | 2023/09/27 10:27:16 | NIST 2702 | NIST 2702 | 0 | 0 | 934 | 6.2467 | 5321 | -24.6504 | QC OK | 7.401 | 0.018387 | 0.22323 | 0.24843 | 3.0163 | 1.3124 | 18.5873 | 12.1411 | 14.163 | QC OK |  | N | N
Acquisition-0170.cf | 2023/09/27 12:42:33 | NIST 2702 | NIST 2702 | 0 | 0 | 561 | 5.9061 | 3268 | -24.7404 | QC OK | 4.28 | 0.01038 | 0.13397 | 0.24251 | 3.1301 | 0.74087 | 11.1547 | 12.9069 | 15.0562 | QC OK |  | N | N
Acquisition-0171.cf | 2023/09/27 12:49:40 | NIST 2702 | NIST 2702 | 0 | 0 | 861 | 5.9738 | 5226 | -24.5519 | QC OK | 6.791 | 0.016745 | 0.21585 | 0.24657 | 3.1785 | 1.1952 | 17.9725 | 12.8906 | 15.0373 | QC OK |  | N | N
Acquisition-0172.cf | 2023/09/27 12:56:48 | NIST 2702 | NIST 2702 | 0 | 0 | 780 | 6.0266 | 4530 | -24.7054 | QC OK | 6.13 | 0.015096 | 0.18771 | 0.24627 | 3.0621 | 1.0775 | 15.6294 | 12.4343 | 14.505 | QC OK |  | N | N
Acquisition-0182.cf | 2023/09/28 10:17:13 | NIST 2702 | NIST 2702 | 0 | 0 | 450 | 6.0005 | 2497 | -24.6655 | QC OK | 3.359 | 0.0085672 | 0.10214 | 0.25505 | 3.0409 | 0.61151 | 8.5049 | 11.9225 | 13.908 | QC OK |  | N | N
Acquisition-0183.cf | 2023/09/28 10:24:20 | NIST 2702 | NIST 2702 | 0 | 0 | 842 | 5.9908 | 4739 | -24.564 | QC OK | 6.616 | 0.016992 | 0.19983 | 0.25684 | 3.0204 | 1.2129 | 16.6388 | 11.7602 | 13.7186 | QC OK |  | N | N
Acquisition-0202.cf | 2023/09/28 12:39:33 | NIST 2702 | NIST 2702 | 0 | 0 | 1233 | 6.069 | 7241 | -24.7246 | QC OK | 9.791 | 0.025237 | 0.29844 | 0.25775 | 3.0481 | 1.8013 | 24.8491 | 11.8256 | 13.7949 | QC OK |  | N | N
Acquisition-0203.cf | 2023/09/28 12:46:40 | NIST 2702 | NIST 2702 | 0 | 0 | 791 | 5.9469 | 4967 | -24.8198 | QC OK | 6.117 | 0.015872 | 0.20422 | 0.25947 | 3.3386 | 1.1329 | 17.0046 | 12.867 | 15.0097 | QC OK |  | N | N
Acquisition-0204.cf | 2023/09/28 12:53:46 | NIST 2702 | NIST 2702 | 0 | 0 | 646 | 5.948 | 3872 | -24.8625 | QC OK | 4.978 | 0.012774 | 0.1579 | 0.2566 | 3.172 | 0.91175 | 13.1475 | 12.3616 | 14.4201 | QC OK |  | N | N
Acquisition-0113.cf | 2023/10/03 10:14:55 | NIST 2702 | NIST 2702 | 0 | 0 | 835 | 6.0868 | 5003 | -24.4885 | QC OK | 6.475 | 0.016442 | 0.19824 | 0.25393 | 3.0617 | 1.1736 | 16.5065 | 12.0571 | 14.0649 | QC OK |  | N | N
Acquisition-0114.cf | 2023/10/03 10:22:02 | NIST 2702 | NIST 2702 | 0 | 0 | 539 | 6.15 | 3161 | -24.5331 | QC OK | 4.065 | 0.010145 | 0.12405 | 0.24958 | 3.0516 | 0.72415 | 10.3285 | 12.2269 | 14.2631 | QC OK |  | N | N
Acquisition-0133.cf | 2023/10/03 12:37:20 | NIST 2702 | NIST 2702 | 0 | 0 | 620 | 5.9547 | 3678 | -24.1234 | QC OK | 4.674 | 0.011961 | 0.14396 | 0.2559 | 3.08 | 0.85373 | 11.9866 | 12.036 | 14.0403 | QC OK |  | N | N
Acquisition-0134.cf | 2023/10/03 12:44:28 | NIST 2702 | NIST 2702 | 0 | 0 | 1617 | 5.9117 | 9748 | -24.2066 | QC OK | 12.879 | 0.033338 | 0.39172 | 0.25885 | 3.0416 | 2.3796 | 32.6164 | 11.7502 | 13.7069 | QC OK |  | N | N
Acquisition-0135.cf | 2023/10/03 12:51:34 | NIST 2702 | NIST 2702 | 0 | 0 | 937 | 5.9966 | 5821 | -24.1714 | QC OK | 7.286 | 0.018848 | 0.23206 | 0.25869 | 3.185 | 1.3453 | 19.3224 | 12.3122 | 14.3626 | QC OK |  | N | N

## SAMPLES
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 (mV) | d15N NIST 8573/8574 corrected (AT-Air) (permil) | Ampl 44 (mV) | d13C NIST 8573/8574 corrected (VPDB) (per mil) | IRMS comment | mass (mg) | N (mg) | C (mg) | N% | C% | N (µmol) | C (µmol) | C:N (mass) | C:N (molar) | EA comment
Acquisition-0120.cf | 2023/09/18 09:50:29 | WS23-139-11MC-3 37.5-40cm | WS23-139-11MC-3 37.5-40cm | 0 | 0 | 924 | 5.2268 | 2902 | -24.1281 | 13C QC OK; DO NOT USE 15N | 16.593 | 0.019166 | 0.12007 | 0.11551 | 0.72361 | 1.368 | 9.9974 | 6.2646 | 7.3078 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0121.cf | 2023/09/18 09:57:37 | WS23-139-10MC-3 15-17.5cm | WS23-139-10MC-3 15-17.5cm | 0 | 0 | 890 | 4.4576 | 2955 | -25.9318 | 13C QC OK; DO NOT USE 15N | 16.358 | 0.018166 | 0.12094 | 0.11105 | 0.73934 | 1.2966 | 10.07 | 6.6576 | 7.7663 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0122.cf | 2023/09/18 10:04:43 | WS23-139-10MC-3 12.5-15cm | WS23-139-10MC-3 12.5-15cm | 0 | 0 | 879 | 4.7628 | 2881 | -25.885 | 13C QC OK; DO NOT USE 15N | 16.17 | 0.017942 | 0.11795 | 0.11096 | 0.72942 | 1.2806 | 9.8207 | 6.5738 | 7.6686 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0123.cf | 2023/09/18 10:11:50 | WS23-139-11MC-3 15-17.5cm | WS23-139-11MC-3 15-17.5cm | 0 | 0 | 1107 | 5.0914 | 3333 | -23.3036 | 13C QC OK; DO NOT USE 15N | 17.567 | 0.022885 | 0.1358 | 0.13028 | 0.77302 | 1.6335 | 11.307 | 5.9338 | 6.9219 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0124.cf | 2023/09/18 10:18:57 | WS23-139-11MC-3 17.5-20cm | WS23-139-11MC-3 17.5-20cm | 0 | 0 | 1092 | 5.0434 | 3245 | -23.4538 | 13C QC OK; DO NOT USE 15N | 17.978 | 0.022616 | 0.13247 | 0.1258 | 0.73687 | 1.6143 | 11.0303 | 5.8575 | 6.833 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0125.cf | 2023/09/18 10:26:04 | WS23-139-10MC-3 30-32.5cm | WS23-139-10MC-3 30-32.5cm | 0 | 0 | 949 | 4.6771 | 3137 | -25.9889 | 13C QC OK; DO NOT USE 15N | 17.341 | 0.019489 | 0.12834 | 0.11239 | 0.74009 | 1.3911 | 10.686 | 6.5851 | 7.6817 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0126.cf | 2023/09/18 10:33:11 | WS23-139-10MC-3 40-42.5cm | WS23-139-10MC-3 40-42.5cm | 0 | 0 | 926 | 4.8872 | 3117 | -26.0041 | 13C QC OK; DO NOT USE 15N | 16.89 | 0.018799 | 0.12669 | 0.1113 | 0.75008 | 1.3418 | 10.5487 | 6.7391 | 7.8614 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0127.cf | 2023/09/18 10:40:17 | WS23-139-10MC-3 27.5-30cm | WS23-139-10MC-3 27.5-30cm | 0 | 0 | 921 | 4.7354 | 3064 | -25.9432 | 13C QC OK; DO NOT USE 15N | 16.929 | 0.018884 | 0.1251 | 0.11155 | 0.73896 | 1.3479 | 10.4161 | 6.6244 | 7.7276 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0128.cf | 2023/09/18 10:47:23 | WS23-139-10MC-3 45-47.5cm | WS23-139-10MC-3 45-47.5cm | 0 | 0 | 892 | 4.7226 | 3004 | -25.9228 | 13C QC OK; DO NOT USE 15N | 16.533 | 0.018169 | 0.12253 | 0.1099 | 0.74113 | 1.2969 | 10.2024 | 6.7439 | 7.8669 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0129.cf | 2023/09/18 10:54:29 | WS23-139-10MC-3 42.5-45cm | WS23-139-10MC-3 42.5-45cm | 0 | 0 | 873 | 4.8835 | 2933 | -26.0058 | 13C QC OK; DO NOT USE 15N | 16.048 | 0.017737 | 0.11927 | 0.11053 | 0.74323 | 1.266 | 9.9312 | 6.7245 | 7.8443 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0130.cf | 2023/09/18 11:01:35 | WS23-139-11MC-3 10-12.5cm | WS23-139-11MC-3 10-12.5cm | 0 | 0 | 1030 | 5.1948 | 3110 | -23.2963 | 13C QC OK; DO NOT USE 15N | 15.717 | 0.020942 | 0.12563 | 0.13325 | 0.79934 | 1.4948 | 10.4606 | 5.999 | 6.9979 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0131.cf | 2023/09/18 11:08:42 | WS23-139-10MC-3 35-37.5cm | WS23-139-10MC-3 35-37.5cm | 0 | 0 | 884 | 4.8156 | 2994 | -26.0475 | 13C QC OK; DO NOT USE 15N | 15.915 | 0.017863 | 0.12163 | 0.11224 | 0.76425 | 1.275 | 10.1275 | 6.809 | 7.9429 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0132.cf | 2023/09/18 11:15:49 | WS23-139-11MC-3 30-32.5cm | WS23-139-11MC-3 30-32.5cm | 0 | 0 | 998 | 4.9762 | 3036 | -23.7428 | 13C QC OK; DO NOT USE 15N | 16.512 | 0.020317 | 0.12288 | 0.12304 | 0.74418 | 1.4502 | 10.2314 | 6.0481 | 7.0552 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0133.cf | 2023/09/18 11:22:56 | WS23-139-11MC-3 12.5-15cm | WS23-139-11MC-3 12.5-15cm | 0 | 0 | 1142 | 5.062 | 3472 | -23.2309 | 13C QC OK; DO NOT USE 15N | 17.732 | 0.023358 | 0.14016 | 0.13173 | 0.79041 | 1.6673 | 11.6699 | 6.0002 | 6.9994 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0134.cf | 2023/09/18 11:30:02 | WS23-139-11MC-3 27.5-30cm | WS23-139-11MC-3 27.5-30cm | 0 | 0 | 995 | 4.9192 | 3069 | -23.8456 | 13C QC OK; DO NOT USE 15N | 16.537 | 0.020286 | 0.12404 | 0.12267 | 0.75007 | 1.448 | 10.328 | 6.1144 | 7.1326 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0135.cf | 2023/09/18 11:37:09 | WS23-139-11MC-3 22.5-25cm | WS23-139-11MC-3 22.5-25cm | 0 | 0 | 1100 | 4.8904 | 3433 | -23.7675 | 13C QC OK; DO NOT USE 15N | 18.011 | 0.022525 | 0.1385 | 0.12506 | 0.76899 | 1.6078 | 11.5322 | 6.1488 | 7.1728 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0136.cf | 2023/09/18 11:44:19 | WS23-139-11MC-3 5-7.5cm | WS23-139-11MC-3 5-7.5cm | 0 | 0 | 1054 | 5.0202 | 3197 | -23.3033 | 13C QC OK; DO NOT USE 15N | 15.957 | 0.021403 | 0.12885 | 0.13413 | 0.80747 | 1.5277 | 10.7284 | 6.0202 | 7.0227 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0137.cf | 2023/09/18 11:51:26 | WS23-139-11MC-3 35-37.5cm | WS23-139-11MC-3 35-37.5cm | 0 | 0 | 989 | 5.0244 | 3058 | -23.8424 | 13C QC OK; DO NOT USE 15N | 17.391 | 0.020182 | 0.1231 | 0.11605 | 0.70783 | 1.4405 | 10.2496 | 6.0995 | 7.1152 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0152.cf | 2023/09/27 10:34:22 | WS23-139-10MC-3 37.5-40cm | WS23-139-10MC-3 37.5-40cm | 0 | 0 | 921 | 4.4286 | 3155 | -26.0594 | 13C QC OK; DO NOT USE 15N | 16.74 | 0.01829 | 0.12956 | 0.10926 | 0.77395 | 1.3055 | 10.7876 | 7.0835 | 8.2631 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0153.cf | 2023/09/27 10:41:30 | WS23-139-10MC-3 47.5-50cm | WS23-139-10MC-3 47.5-50cm | 0 | 0 | 936 | 4.6415 | 3093 | -25.6954 | 13C QC OK; DO NOT USE 15N | 17.742 | 0.018747 | 0.12737 | 0.10566 | 0.71791 | 1.3381 | 10.6054 | 6.7943 | 7.9258 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0154.cf | 2023/09/27 10:48:37 | WS23-139-11MC-3 32.5-35cm | WS23-139-11MC-3 32.5-35cm | 0 | 0 | 927 | 5.064 | 2702 | -23.5805 | 13C QC OK; DO NOT USE 15N | 15.88 | 0.018345 | 0.11004 | 0.11552 | 0.69295 | 1.3094 | 9.1624 | 5.9983 | 6.9972 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0155.cf | 2023/09/27 10:55:44 | WS23-139-11MC-3 25-27.5cm | WS23-139-11MC-3 25-27.5cm | 0 | 0 | 1011 | 4.9961 | 3128 | -23.7862 | 13C QC OK; DO NOT USE 15N | 16.993 | 0.020377 | 0.12853 | 0.11992 | 0.75638 | 1.4545 | 10.7021 | 6.3076 | 7.358 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0156.cf | 2023/09/27 11:02:51 | WS23-139-11MC-3 7.5-10cm | WS23-139-11MC-3 7.5-10cm | 0 | 0 | 1114 | 4.653 | 3355 | -23.2504 | 13C QC OK; DO NOT USE 15N | 17.201 | 0.022369 | 0.13812 | 0.13005 | 0.80295 | 1.5967 | 11.5001 | 6.1744 | 7.2026 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0157.cf | 2023/09/27 11:09:59 | WS23-139-10MC-3 32.5-35cm | WS23-139-10MC-3 32.5-35cm | 0 | 0 | 967 | 4.5913 | 3262 | -25.9362 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.019154 | 0.1325 | 0.11117 | 0.76906 | 1.3672 | 11.0326 | 6.9178 | 8.0698 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0158.cf | 2023/09/27 11:17:06 | WS23-139-10MC-3 10-12.5cm | WS23-139-10MC-3 10-12.5cm | 0 | 0 | 992 | 4.7649 | 3484 | -25.9944 | 13C QC OK; DO NOT USE 15N | 17.631 | 0.019684 | 0.14166 | 0.11164 | 0.80349 | 1.405 | 11.7954 | 7.1969 | 8.3954 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0159.cf | 2023/09/27 11:24:13 | WS23-139-11MC-3 2.5-5cm | WS23-139-11MC-3 2.5-5cm | 0 | 0 | 1072 | 5.2304 | 3232 | -23.2404 | 13C QC OK; DO NOT USE 15N | 16.03 | 0.021205 | 0.13131 | 0.13228 | 0.81915 | 1.5136 | 10.9333 | 6.1924 | 7.2236 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0160.cf | 2023/09/27 11:31:20 | WS23-139-11MC-3 0-2.5cm | WS23-139-11MC-3 0-2.5cm | 0 | 0 | 1089 | 5.0481 | 3276 | -23.1954 | 13C QC OK; DO NOT USE 15N | 16.223 | 0.021633 | 0.13351 | 0.13335 | 0.82294 | 1.5441 | 11.1162 | 6.1713 | 7.1989 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0161.cf | 2023/09/27 11:38:27 | WS23-139-11MC-3 20-22.5cm | WS23-139-11MC-3 20-22.5cm | 0 | 0 | 1015 | 5.1612 | 3103 | -23.5766 | 13C QC OK; DO NOT USE 15N | 16.656 | 0.020157 | 0.12591 | 0.12102 | 0.75593 | 1.4388 | 10.4835 | 6.2463 | 7.2865 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0162.cf | 2023/09/27 11:45:34 | WS23-139-16MC-3 15-17.5cm | WS23-139-16MC-3 15-17.5cm | 0 | 0 | 937 | 5.2242 | 2843 | -23.5833 | 13C QC OK; DO NOT USE 15N | 17.495 | 0.018346 | 0.11556 | 0.10487 | 0.66055 | 1.3095 | 9.6222 | 6.299 | 7.3479 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0163.cf | 2023/09/27 11:52:42 | WS23-139-18MC-3 7.5-10cm | WS23-139-18MC-3 7.5-10cm | 0 | 0 | 1013 | 5.0938 | 3323 | -23.3657 | 13C QC OK; DO NOT USE 15N | 17.172 | 0.020137 | 0.13561 | 0.11727 | 0.78969 | 1.4373 | 11.291 | 6.7342 | 7.8556 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0164.cf | 2023/09/27 11:59:49 | WS23-139-18MC-3 12.5-15cm | WS23-139-18MC-3 12.5-15cm | 0 | 0 | 1017 | 5.2125 | 3217 | -23.2717 | 13C QC OK; DO NOT USE 15N | 18.262 | 0.020279 | 0.13206 | 0.11104 | 0.72316 | 1.4475 | 10.9962 | 6.5124 | 7.5969 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0165.cf | 2023/09/27 12:06:56 | WS23-139-01MC-3 5-7.5cm | WS23-139-01MC-3 5-7.5cm | 0 | 0 | 364 | 4.3964 | 1778 | -27.7181 | 13C QC OK; DO NOT USE 15N | 16.748 | 0.0060456 | 0.072292 | 0.036097 | 0.43165 | 0.43152 | 6.0194 | 11.9579 | 13.9492 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0166.cf | 2023/09/27 12:14:04 | WS23-139-01MC-3 22.5-25cm | WS23-139-01MC-3 22.5-25cm | 0 | 0 | 302 | 4.0424 | 684 | -25.7423 | 13C QC OK; DO NOT USE 15N | 17.698 | 0.0046911 | 0.02699 | 0.026507 | 0.1525 | 0.33484 | 2.2473 | 5.7533 | 6.7114 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0167.cf | 2023/09/27 12:21:11 | WS23-139-16MC-3 32.5-35cm | WS23-139-16MC-3 32.5-35cm | 0 | 0 | 996 | 5.1852 | 3058 | -23.6327 | 13C QC OK; DO NOT USE 15N | 18.655 | 0.019621 | 0.12369 | 0.10518 | 0.66305 | 1.4005 | 10.2991 | 6.304 | 7.3537 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0168.cf | 2023/09/27 12:28:18 | WS23-139-01MC-3 12.5-15cm | WS23-139-01MC-3 12.5-15cm | 0 | 0 | 296 | 3.6666 | 588 | -25.6388 | 13C QC OK; DO NOT USE 15N | 17.957 | 0.0045648 | 0.023059 | 0.02542 | 0.12841 | 0.32582 | 1.92 | 5.0516 | 5.8929 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0169.cf | 2023/09/27 12:35:26 | WS23-139-10MC-3 2.5-5.0cm | WS23-139-10MC-3 2.5-5.0cm | 0 | 0 | 1084 | 4.7995 | 3819 | -25.836 | 13C QC OK; DO NOT USE 15N | 18.758 | 0.021591 | 0.15526 | 0.1151 | 0.8277 | 1.5411 | 12.9276 | 7.191 | 8.3885 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0184.cf | 2023/09/28 10:31:28 | WS23-139-10MC-3 22.5-25cm | WS23-139-10MC-3 22.5-25cm | 0 | 0 | 1014 | 5.099 | 3390 | -25.9015 | 13C QC OK; DO NOT USE 15N | 18.597 | 0.021035 | 0.13884 | 0.11311 | 0.74658 | 1.5014 | 11.5604 | 6.6005 | 7.6997 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0185.cf | 2023/09/28 10:38:34 | WS23-139-18MC-3 15-17.5cm | WS23-139-18MC-3 15-17.5cm | 0 | 0 | 1100 | 4.9445 | 3572 | -23.2996 | 13C QC OK; DO NOT USE 15N | 18.433 | 0.022651 | 0.14555 | 0.12288 | 0.78963 | 1.6168 | 12.1193 | 6.4258 | 7.4959 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0186.cf | 2023/09/28 10:45:42 | WS23-139-01MC-3 27.5-30cm | WS23-139-01MC-3 27.5-30cm | 0 | 0 | 300 | 4.3353 | 712 | -25.9516 | 13C QC OK; DO NOT USE 15N | 17.828 | 0.0053308 | 0.028472 | 0.029901 | 0.1597 | 0.3805 | 2.3707 | 5.341 | 6.2304 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0187.cf | 2023/09/28 10:52:47 | WS23-139-16MC-3 35-36cm | WS23-139-16MC-3 35-36cm | 0 | 0 | 998 | 5.0488 | 3020 | -23.7158 | 13C QC OK; DO NOT USE 15N | 18.705 | 0.020205 | 0.12118 | 0.10802 | 0.64785 | 1.4422 | 10.09 | 5.9975 | 6.9963 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0188.cf | 2023/09/28 10:59:55 | WS23-139-01MC-3 0-2.5cm | WS23-139-01MC-3 0-2.5cm | 0 | 0 | 385 | 4.0457 | 1692 | -27.2857 | 13C QC OK; DO NOT USE 15N | 17.14 | 0.007201 | 0.068228 | 0.042013 | 0.39806 | 0.51399 | 5.6809 | 9.4747 | 11.0525 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0189.cf | 2023/09/28 11:07:02 | WS23-139-16MC-3 22.5-25cm | WS23-139-16MC-3 22.5-25cm | 0 | 0 | 887 | 5.026 | 2713 | -23.7688 | 13C QC OK; DO NOT USE 15N | 16.649 | 0.018007 | 0.10913 | 0.10816 | 0.65548 | 1.2853 | 9.0867 | 6.0605 | 7.0697 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0190.cf | 2023/09/28 11:14:09 | WS23-139-18MC-3 2.5-5cm | WS23-139-18MC-3 2.5-5cm | 0 | 0 | 1129 | 5.2451 | 3779 | -23.0664 | 13C QC OK; DO NOT USE 15N | 16.532 | 0.023288 | 0.15285 | 0.14086 | 0.92457 | 1.6622 | 12.7269 | 6.5636 | 7.6566 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0191.cf | 2023/09/28 11:21:17 | WS23-139-10MC-3 20-22.5cm | WS23-139-10MC-3 20-22.5cm | 0 | 0 | 984 | 4.6768 | 3256 | -25.8781 | 13C QC OK; DO NOT USE 15N | 17.743 | 0.020275 | 0.13143 | 0.11427 | 0.74075 | 1.4472 | 10.9435 | 6.4825 | 7.562 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0192.cf | 2023/09/28 11:28:24 | WS23-139-01MC-3 2.5-5cm | WS23-139-01MC-3 2.5-5cm | 0 | 0 | 226 | 4.2666 | 646 | -26.8383 | 13C QC OK; DO NOT USE 15N | 16.612 | 0.0036941 | 0.025452 | 0.022237 | 0.15322 | 0.26367 | 2.1193 | 6.89 | 8.0374 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0193.cf | 2023/09/28 11:35:29 | WS23-139-01MC-3 25-27.5cm | WS23-139-01MC-3 25-27.5cm | 0 | 0 | 348 | 3.9236 | 815 | -26.0453 | 13C QC OK; DO NOT USE 15N | 18.559 | 0.0063625 | 0.032427 | 0.034282 | 0.17472 | 0.45414 | 2.7 | 5.0966 | 5.9454 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0194.cf | 2023/09/28 11:42:35 | WS23-139-10MC-3 17.5-20cm | WS23-139-10MC-3 17.5-20cm | 0 | 0 | 993 | 4.6877 | 3310 | -25.834 | 13C QC OK; DO NOT USE 15N | 17.982 | 0.020398 | 0.13342 | 0.11344 | 0.74198 | 1.456 | 11.1092 | 6.5409 | 7.6301 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0195.cf | 2023/09/28 11:49:42 | WS23-139-18MC-3 17.5-20cm | WS23-139-18MC-3 17.5-20cm | 0 | 0 | 915 | 5.0653 | 2842 | -23.3299 | 13C QC OK; DO NOT USE 15N | 17.305 | 0.018686 | 0.11598 | 0.10798 | 0.67018 | 1.3338 | 9.6565 | 6.2064 | 7.24 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0196.cf | 2023/09/28 11:56:49 | WS23-139-10MC-3 0-2.5cm | WS23-139-10MC-3 0-2.5cm | 0 | 0 | 1046 | 4.7378 | 3455 | -25.5252 | 13C QC OK; DO NOT USE 15N | 17.36 | 0.021343 | 0.13896 | 0.12295 | 0.80045 | 1.5234 | 11.5702 | 6.5106 | 7.5948 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0197.cf | 2023/09/28 12:03:57 | WS23-139-16MC-3 25-27.5cm | WS23-139-16MC-3 25-27.5cm | 0 | 0 | 920 | 5.015 | 2800 | -23.7736 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.018699 | 0.11345 | 0.10853 | 0.65847 | 1.3347 | 9.4462 | 6.0672 | 7.0776 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0198.cf | 2023/09/28 12:11:04 | WS23-139-10MC-3 7.5-10cm | WS23-139-10MC-3 7.5-10cm | 0 | 0 | 1037 | 4.6534 | 3494 | -25.8789 | 13C QC OK; DO NOT USE 15N | 18.667 | 0.02124 | 0.14043 | 0.11378 | 0.75229 | 1.5161 | 11.6927 | 6.6115 | 7.7125 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0199.cf | 2023/09/28 12:18:11 | WS23-139-01MC-3 10-12.5cm | WS23-139-01MC-3 10-12.5cm | 0 | 0 | 405 | 4.4545 | 1279 | -26.5615 | 13C QC OK; DO NOT USE 15N | 18.039 | 0.0075749 | 0.051233 | 0.041992 | 0.28401 | 0.54068 | 4.2659 | 6.7636 | 7.8899 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0200.cf | 2023/09/28 12:25:18 | WS23-139-10MC-3 25-27.5cm | WS23-139-10MC-3 25-27.5cm | 0 | 0 | 957 | 4.6191 | 3173 | -25.9636 | 13C QC OK; DO NOT USE 15N | 17.002 | 0.01948 | 0.12782 | 0.11457 | 0.75181 | 1.3904 | 10.6431 | 6.5618 | 7.6545 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0201.cf | 2023/09/28 12:32:26 | WS23-139-16MC-3 30-32.5cm | WS23-139-16MC-3 30-32.5cm | 0 | 0 | 951 | 5.3283 | 2899 | -23.7796 | 13C QC OK; DO NOT USE 15N | 17.75 | 0.01934 | 0.11716 | 0.10896 | 0.66005 | 1.3804 | 9.7551 | 6.0579 | 7.0668 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0115.cf | 2023/10/03 10:29:10 | WS23-139-10MC-3 5-7.5cm | WS23-139-10MC-3 5-7.5cm | 0 | 0 | 936 | 4.6265 | 3230 | -25.8169 | 13C QC OK; DO NOT USE 15N | 16.815 | 0.018881 | 0.12502 | 0.11229 | 0.7435 | 1.3477 | 10.4096 | 6.6212 | 7.7239 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0116.cf | 2023/10/03 10:36:17 | WS23-139-18MC-3 0-2.5cm | WS23-139-18MC-3 0-2.5cm | 0 | 0 | 936 | 5.0125 | 3043 | -22.99 | 13C QC OK; DO NOT USE 15N | 17.304 | 0.018746 | 0.11734 | 0.10833 | 0.67809 | 1.338 | 9.7699 | 6.2593 | 7.3017 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0117.cf | 2023/10/03 10:43:24 | WS23-139-01MC-3 17.5-20cm | WS23-139-01MC-3 17.5-20cm | 0 | 0 | 288 | 4.7768 | 603 | -25.6501 | 13C QC OK; DO NOT USE 15N | 16.608 | 0.0047092 | 0.022542 | 0.028355 | 0.13573 | 0.33613 | 1.877 | 4.7869 | 5.5841 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0118.cf | 2023/10/03 10:50:31 | WS23-139-01MC-3 15-17.5cm | WS23-139-01MC-3 15-17.5cm | 0 | 0 | 346 | 4.302 | 801 | -25.4707 | 13C QC OK; DO NOT USE 15N | 18.809 | 0.005988 | 0.03044 | 0.031836 | 0.16184 | 0.42741 | 2.5346 | 5.0835 | 5.93 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0119.cf | 2023/10/03 10:57:39 | WS23-139-16MC-3 12.5-15cm | WS23-139-16MC-3 12.5-15cm | 0 | 0 | 885 | 5.0606 | 2788 | -23.5386 | 13C QC OK; DO NOT USE 15N | 16.565 | 0.017552 | 0.10758 | 0.10596 | 0.64946 | 1.2528 | 8.9578 | 6.1294 | 7.1501 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0120.cf | 2023/10/03 11:04:46 | WS23-139-16MC-3 20-22.5cm | WS23-139-16MC-3 20-22.5cm | 0 | 0 | 943 | 5.1821 | 2991 | -23.5546 | 13C QC OK; DO NOT USE 15N | 17.78 | 0.018861 | 0.11531 | 0.10608 | 0.64855 | 1.3462 | 9.6014 | 6.1139 | 7.132 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0121.cf | 2023/10/03 11:11:54 | WS23-139-16MC-3 17.5-20cm | WS23-139-16MC-3 17.5-20cm | 0 | 0 | 1003 | 5.1417 | 3189 | -23.4979 | 13C QC OK; DO NOT USE 15N | 18.789 | 0.019999 | 0.12264 | 0.10644 | 0.65275 | 1.4275 | 10.2119 | 6.1324 | 7.1536 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0122.cf | 2023/10/03 11:19:01 | WS23-139-18MC-3 20-22.5cm | WS23-139-18MC-3 20-22.5cm | 0 | 0 | 967 | 4.8509 | 3227 | -23.0488 | 13C QC OK; DO NOT USE 15N | 18.639 | 0.019427 | 0.12544 | 0.10423 | 0.67302 | 1.3867 | 10.445 | 6.4572 | 7.5325 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0123.cf | 2023/10/03 11:26:08 | WS23-139-16MC-3 10-12.5cm | WS23-139-16MC-3 10-12.5cm | 0 | 0 | 900 | 5.025 | 2852 | -23.5544 | 13C QC OK; DO NOT USE 15N | 16.701 | 0.017805 | 0.10965 | 0.10661 | 0.65656 | 1.2709 | 9.13 | 6.1585 | 7.1841 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0124.cf | 2023/10/03 11:33:16 | WS23-139-16MC-3 27.5-30cm | WS23-139-16MC-3 27.5-30cm | 0 | 0 | 1004 | 4.9667 | 3171 | -23.4367 | 13C QC OK; DO NOT USE 15N | 18.643 | 0.02009 | 0.12199 | 0.10776 | 0.65436 | 1.434 | 10.1575 | 6.0722 | 7.0834 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0125.cf | 2023/10/03 11:40:23 | WS23-139-17MC-3 45-47.5cm | WS23-139-17MC-3 45-47.5cm | 0 | 0 | 1550 | 4.9085 | 9124 | -26.871 | 13C QC OK; DO NOT USE 15N | 18.472 | 0.031744 | 0.36142 | 0.17185 | 1.9566 | 2.2658 | 30.093 | 11.3854 | 13.2814 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0126.cf | 2023/10/03 11:47:30 | WS23-139-01MC-3 7.5-10cm | WS23-139-01MC-3 7.5-10cm | 0 | 0 | 642 | 4.2587 | 3112 | -26.7986 | 13C QC OK; DO NOT USE 15N | 18.293 | 0.012449 | 0.12008 | 0.068054 | 0.65645 | 0.88859 | 9.9986 | 9.646 | 11.2523 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0127.cf | 2023/10/03 11:54:37 | WS23-139-01MC-3 20-22.5cm | WS23-139-01MC-3 20-22.5cm | 0 | 0 | 267 | 3.8766 | 575 | -25.1163 | 13C QC OK; DO NOT USE 15N | 16.12 | 0.0042932 | 0.021494 | 0.026633 | 0.13334 | 0.30644 | 1.7897 | 5.0065 | 5.8402 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0128.cf | 2023/10/03 12:01:44 | WS23-139-18MC-3 10-12.5cm | WS23-139-18MC-3 10-12.5cm | 0 | 0 | 824 | 4.7839 | 2611 | -23.0869 | 13C QC OK; DO NOT USE 15N | 17.429 | 0.016437 | 0.10152 | 0.094311 | 0.58248 | 1.1733 | 8.453 | 6.1762 | 7.2047 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0129.cf | 2023/10/03 12:08:52 | WS23-139-17MC-3 47.5-48.75cm | WS23-139-17MC-3 47.5-48.75cm | 0 | 0 | 1294 | 5.0464 | 6054 | -26.1607 | 13C QC OK; DO NOT USE 15N | 16.697 | 0.026758 | 0.24203 | 0.16026 | 1.4496 | 1.9099 | 20.1527 | 9.0453 | 10.5515 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0130.cf | 2023/10/03 12:15:59 | WS23-139-18MC-3 5-7.5cm | WS23-139-18MC-3 5-7.5cm | 0 | 0 | 1028 | 5.0278 | 3550 | -22.821 | 13C QC OK; DO NOT USE 15N | 16.244 | 0.020788 | 0.13736 | 0.12797 | 0.84558 | 1.4838 | 11.4368 | 6.6075 | 7.7079 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0131.cf | 2023/10/03 12:23:05 | WS23-139-05MC-3 25-27.5cm | WS23-139-05MC-3 25-27.5cm | 0 | 0 | 637 | 4.4392 | 2322 | -26.367 | 13C QC OK; DO NOT USE 15N | 17.952 | 0.012453 | 0.089876 | 0.069366 | 0.50064 | 0.88883 | 7.4834 | 7.2174 | 8.4193 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0132.cf | 2023/10/03 12:30:13 | WS23-139-05MC-3 52.5-55cm | WS23-139-05MC-3 52.5-55cm | 0 | 0 | 776 | 4.5607 | 2693 | -26.3466 | 13C QC OK; DO NOT USE 15N | 19.005 | 0.015525 | 0.1048 | 0.08169 | 0.55143 | 1.1081 | 8.726 | 6.7503 | 7.8745 | C QC OK; N and C:N use discretion, these are acidified samples | N | N

## summary stats
ANALYTICAL UNCERTAINTY - estimates expressed as ±1 standard deviation of replicate measurements of "working" SRM
SRM = | NIST 2702
n= | 20
ELEMENTAL COMPOSITION - calibrated to NIST 8573 & NIST 8574 SRM
| STDEV | RSD
± | 0.0054961 | 2.174% | %N
± | 0.08765 | 2.83% | %C
± | 0.38349 | 3.129% | C:N (mass)
± | 0.44735 | 3.129% | C:N (molar)
ISOTOPIC COMPOSITION - calibrated/normalized to NIST 8573 & NIST 8574 SRM
± | 0.10183 |  | d15N (reported permil (‰) relative to AT-Air)
± | 0.20325 |  | d13C (reported permil (‰) relative to VPDB)

## RPT
ANALYTICAL UNCERTAINTY - estimates expressed as ±1 standard deviation of replicate measurements of "working" SRM | [10 empty] | REFERENCE MATERIAL SPECIFICATIONS
SRM = | NIST 2702 | [9 empty] | RM | Material | d15N (‰, AT-Air) | ±1sd (‰) | d13C (‰, VPDB) | ±1sd (‰) | N% | C% | C:N (mass) | C:N (molar) | Specifications based on
n= | 20 | [9 empty] | USGS 40 | L-glutamic acid | -4.52 | 0.12 | -26.39 | 0.09 | 9.52 | 40.81 | 4.2868 | 5.0021 | USGS (NIST) specifications - primary std for AT-Air and VPDB scale normalization
[11 empty] | USGS 41a | L-glutamic acid | 47.55 | 0.15 | 36.55 | 0.08 | 9.52 | 40.81 | 4.2868 | 5.0021 | USGS (NIST) specifications - primary std for AT-Air and VPDB scale normalization
ELEMENTAL COMPOSITION - calibrated to NIST 8573 & NIST 8574 SRM | [10 empty] | NIST 2702 | sediment (marine) | 5.95 | 0.25 | -24.49 | 0.09 | 0.26 | 3.11 | 11.86 | 13.84 | Lab Standard NIST 2702 LT tracking.xlsx
| STDEV | RSD
± | 0.0054961 | 0.021742 | %N
± | 0.08765 | 0.0283 | %C
± | 0.38349 | 0.03129 | C:N (mass)
± | 0.44735 | 0.03129 | C:N (molar)
ISOTOPIC COMPOSITION - calibrated/normalized to NIST 8573 & NIST 8574 SRM
± | 0.10183 |  | d15N (reported permil (‰) relative to AT-Air)
± | 0.20325 |  | d13C (reported permil (‰) relative to VPDB)
MEASUREMENTS (n) | [9 empty] | MEASURMENTS PASSING IRMS QC (n) | [9 empty] | MEASURMENTS PASSING EA QC (n)
72 | [9 empty] | 72 | [9 empty] | 72
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 (mV) | d15N (‰, AT-Air) | Ampl 44 (mV) | d13C (‰, VPDB) | IRMS comment | mass (mg) | N (mg) | C (mg) | N% | C% | N (µmol) | C (µmol) | C:N (mass) | C:N (molar) | EA comment
Acquisition-0188.cf | 2023/09/28 10:59:55 | WS23-139-01MC-3 0-2.5cm | WS23-139-01MC-3 0-2.5cm | 0 | 0 | 385 | 4.0457 | 1692 | -27.2857 | 13C QC OK; DO NOT USE 15N | 17.14 | 0.007201 | 0.068228 | 0.042013 | 0.39806 | 0.51399 | 5.6809 | 9.4747 | 11.0525 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0199.cf | 2023/09/28 12:18:11 | WS23-139-01MC-3 10-12.5cm | WS23-139-01MC-3 10-12.5cm | 0 | 0 | 405 | 4.4545 | 1279 | -26.5615 | 13C QC OK; DO NOT USE 15N | 18.039 | 0.0075749 | 0.051233 | 0.041992 | 0.28401 | 0.54068 | 4.2659 | 6.7636 | 7.8899 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0168.cf | 2023/09/27 12:28:18 | WS23-139-01MC-3 12.5-15cm | WS23-139-01MC-3 12.5-15cm | 0 | 0 | 296 | 3.6666 | 588 | -25.6388 | 13C QC OK; DO NOT USE 15N | 17.957 | 0.0045648 | 0.023059 | 0.02542 | 0.12841 | 0.32582 | 1.92 | 5.0516 | 5.8929 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0118.cf | 2023/10/03 10:50:31 | WS23-139-01MC-3 15-17.5cm | WS23-139-01MC-3 15-17.5cm | 0 | 0 | 346 | 4.302 | 801 | -25.4707 | 13C QC OK; DO NOT USE 15N | 18.809 | 0.005988 | 0.03044 | 0.031836 | 0.16184 | 0.42741 | 2.5346 | 5.0835 | 5.93 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0117.cf | 2023/10/03 10:43:24 | WS23-139-01MC-3 17.5-20cm | WS23-139-01MC-3 17.5-20cm | 0 | 0 | 288 | 4.7768 | 603 | -25.6501 | 13C QC OK; DO NOT USE 15N | 16.608 | 0.0047092 | 0.022542 | 0.028355 | 0.13573 | 0.33613 | 1.877 | 4.7869 | 5.5841 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0192.cf | 2023/09/28 11:28:24 | WS23-139-01MC-3 2.5-5cm | WS23-139-01MC-3 2.5-5cm | 0 | 0 | 226 | 4.2666 | 646 | -26.8383 | 13C QC OK; DO NOT USE 15N | 16.612 | 0.0036941 | 0.025452 | 0.022237 | 0.15322 | 0.26367 | 2.1193 | 6.89 | 8.0374 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0127.cf | 2023/10/03 11:54:37 | WS23-139-01MC-3 20-22.5cm | WS23-139-01MC-3 20-22.5cm | 0 | 0 | 267 | 3.8766 | 575 | -25.1163 | 13C QC OK; DO NOT USE 15N | 16.12 | 0.0042932 | 0.021494 | 0.026633 | 0.13334 | 0.30644 | 1.7897 | 5.0065 | 5.8402 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0166.cf | 2023/09/27 12:14:04 | WS23-139-01MC-3 22.5-25cm | WS23-139-01MC-3 22.5-25cm | 0 | 0 | 302 | 4.0424 | 684 | -25.7423 | 13C QC OK; DO NOT USE 15N | 17.698 | 0.0046911 | 0.02699 | 0.026507 | 0.1525 | 0.33484 | 2.2473 | 5.7533 | 6.7114 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0193.cf | 2023/09/28 11:35:29 | WS23-139-01MC-3 25-27.5cm | WS23-139-01MC-3 25-27.5cm | 0 | 0 | 348 | 3.9236 | 815 | -26.0453 | 13C QC OK; DO NOT USE 15N | 18.559 | 0.0063625 | 0.032427 | 0.034282 | 0.17472 | 0.45414 | 2.7 | 5.0966 | 5.9454 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0186.cf | 2023/09/28 10:45:42 | WS23-139-01MC-3 27.5-30cm | WS23-139-01MC-3 27.5-30cm | 0 | 0 | 300 | 4.3353 | 712 | -25.9516 | 13C QC OK; DO NOT USE 15N | 17.828 | 0.0053308 | 0.028472 | 0.029901 | 0.1597 | 0.3805 | 2.3707 | 5.341 | 6.2304 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0165.cf | 2023/09/27 12:06:56 | WS23-139-01MC-3 5-7.5cm | WS23-139-01MC-3 5-7.5cm | 0 | 0 | 364 | 4.3964 | 1778 | -27.7181 | 13C QC OK; DO NOT USE 15N | 16.748 | 0.0060456 | 0.072292 | 0.036097 | 0.43165 | 0.43152 | 6.0194 | 11.9579 | 13.9492 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0126.cf | 2023/10/03 11:47:30 | WS23-139-01MC-3 7.5-10cm | WS23-139-01MC-3 7.5-10cm | 0 | 0 | 642 | 4.2587 | 3112 | -26.7986 | 13C QC OK; DO NOT USE 15N | 18.293 | 0.012449 | 0.12008 | 0.068054 | 0.65645 | 0.88859 | 9.9986 | 9.646 | 11.2523 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ |  | N | N
Acquisition-0131.cf | 2023/10/03 12:23:05 | WS23-139-05MC-3 25-27.5cm | WS23-139-05MC-3 25-27.5cm | 0 | 0 | 637 | 4.4392 | 2322 | -26.367 | 13C QC OK; DO NOT USE 15N | 17.952 | 0.012453 | 0.089876 | 0.069366 | 0.50064 | 0.88883 | 7.4834 | 7.2174 | 8.4193 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0132.cf | 2023/10/03 12:30:13 | WS23-139-05MC-3 52.5-55cm | WS23-139-05MC-3 52.5-55cm | 0 | 0 | 776 | 4.5607 | 2693 | -26.3466 | 13C QC OK; DO NOT USE 15N | 19.005 | 0.015525 | 0.1048 | 0.08169 | 0.55143 | 1.1081 | 8.726 | 6.7503 | 7.8745 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0196.cf | 2023/09/28 11:56:49 | WS23-139-10MC-3 0-2.5cm | WS23-139-10MC-3 0-2.5cm | 0 | 0 | 1046 | 4.7378 | 3455 | -25.5252 | 13C QC OK; DO NOT USE 15N | 17.36 | 0.021343 | 0.13896 | 0.12295 | 0.80045 | 1.5234 | 11.5702 | 6.5106 | 7.5948 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0158.cf | 2023/09/27 11:17:06 | WS23-139-10MC-3 10-12.5cm | WS23-139-10MC-3 10-12.5cm | 0 | 0 | 992 | 4.7649 | 3484 | -25.9944 | 13C QC OK; DO NOT USE 15N | 17.631 | 0.019684 | 0.14166 | 0.11164 | 0.80349 | 1.405 | 11.7954 | 7.1969 | 8.3954 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0122.cf | 2023/09/18 10:04:43 | WS23-139-10MC-3 12.5-15cm | WS23-139-10MC-3 12.5-15cm | 0 | 0 | 879 | 4.7628 | 2881 | -25.885 | 13C QC OK; DO NOT USE 15N | 16.17 | 0.017942 | 0.11795 | 0.11096 | 0.72942 | 1.2806 | 9.8207 | 6.5738 | 7.6686 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0121.cf | 2023/09/18 09:57:37 | WS23-139-10MC-3 15-17.5cm | WS23-139-10MC-3 15-17.5cm | 0 | 0 | 890 | 4.4576 | 2955 | -25.9318 | 13C QC OK; DO NOT USE 15N | 16.358 | 0.018166 | 0.12094 | 0.11105 | 0.73934 | 1.2966 | 10.07 | 6.6576 | 7.7663 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0194.cf | 2023/09/28 11:42:35 | WS23-139-10MC-3 17.5-20cm | WS23-139-10MC-3 17.5-20cm | 0 | 0 | 993 | 4.6877 | 3310 | -25.834 | 13C QC OK; DO NOT USE 15N | 17.982 | 0.020398 | 0.13342 | 0.11344 | 0.74198 | 1.456 | 11.1092 | 6.5409 | 7.6301 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0169.cf | 2023/09/27 12:35:26 | WS23-139-10MC-3 2.5-5.0cm | WS23-139-10MC-3 2.5-5.0cm | 0 | 0 | 1084 | 4.7995 | 3819 | -25.836 | 13C QC OK; DO NOT USE 15N | 18.758 | 0.021591 | 0.15526 | 0.1151 | 0.8277 | 1.5411 | 12.9276 | 7.191 | 8.3885 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0191.cf | 2023/09/28 11:21:17 | WS23-139-10MC-3 20-22.5cm | WS23-139-10MC-3 20-22.5cm | 0 | 0 | 984 | 4.6768 | 3256 | -25.8781 | 13C QC OK; DO NOT USE 15N | 17.743 | 0.020275 | 0.13143 | 0.11427 | 0.74075 | 1.4472 | 10.9435 | 6.4825 | 7.562 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0184.cf | 2023/09/28 10:31:28 | WS23-139-10MC-3 22.5-25cm | WS23-139-10MC-3 22.5-25cm | 0 | 0 | 1014 | 5.099 | 3390 | -25.9015 | 13C QC OK; DO NOT USE 15N | 18.597 | 0.021035 | 0.13884 | 0.11311 | 0.74658 | 1.5014 | 11.5604 | 6.6005 | 7.6997 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0200.cf | 2023/09/28 12:25:18 | WS23-139-10MC-3 25-27.5cm | WS23-139-10MC-3 25-27.5cm | 0 | 0 | 957 | 4.6191 | 3173 | -25.9636 | 13C QC OK; DO NOT USE 15N | 17.002 | 0.01948 | 0.12782 | 0.11457 | 0.75181 | 1.3904 | 10.6431 | 6.5618 | 7.6545 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0127.cf | 2023/09/18 10:40:17 | WS23-139-10MC-3 27.5-30cm | WS23-139-10MC-3 27.5-30cm | 0 | 0 | 921 | 4.7354 | 3064 | -25.9432 | 13C QC OK; DO NOT USE 15N | 16.929 | 0.018884 | 0.1251 | 0.11155 | 0.73896 | 1.3479 | 10.4161 | 6.6244 | 7.7276 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0125.cf | 2023/09/18 10:26:04 | WS23-139-10MC-3 30-32.5cm | WS23-139-10MC-3 30-32.5cm | 0 | 0 | 949 | 4.6771 | 3137 | -25.9889 | 13C QC OK; DO NOT USE 15N | 17.341 | 0.019489 | 0.12834 | 0.11239 | 0.74009 | 1.3911 | 10.686 | 6.5851 | 7.6817 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0157.cf | 2023/09/27 11:09:59 | WS23-139-10MC-3 32.5-35cm | WS23-139-10MC-3 32.5-35cm | 0 | 0 | 967 | 4.5913 | 3262 | -25.9362 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.019154 | 0.1325 | 0.11117 | 0.76906 | 1.3672 | 11.0326 | 6.9178 | 8.0698 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0131.cf | 2023/09/18 11:08:42 | WS23-139-10MC-3 35-37.5cm | WS23-139-10MC-3 35-37.5cm | 0 | 0 | 884 | 4.8156 | 2994 | -26.0475 | 13C QC OK; DO NOT USE 15N | 15.915 | 0.017863 | 0.12163 | 0.11224 | 0.76425 | 1.275 | 10.1275 | 6.809 | 7.9429 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0152.cf | 2023/09/27 10:34:22 | WS23-139-10MC-3 37.5-40cm | WS23-139-10MC-3 37.5-40cm | 0 | 0 | 921 | 4.4286 | 3155 | -26.0594 | 13C QC OK; DO NOT USE 15N | 16.74 | 0.01829 | 0.12956 | 0.10926 | 0.77395 | 1.3055 | 10.7876 | 7.0835 | 8.2631 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0126.cf | 2023/09/18 10:33:11 | WS23-139-10MC-3 40-42.5cm | WS23-139-10MC-3 40-42.5cm | 0 | 0 | 926 | 4.8872 | 3117 | -26.0041 | 13C QC OK; DO NOT USE 15N | 16.89 | 0.018799 | 0.12669 | 0.1113 | 0.75008 | 1.3418 | 10.5487 | 6.7391 | 7.8614 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0129.cf | 2023/09/18 10:54:29 | WS23-139-10MC-3 42.5-45cm | WS23-139-10MC-3 42.5-45cm | 0 | 0 | 873 | 4.8835 | 2933 | -26.0058 | 13C QC OK; DO NOT USE 15N | 16.048 | 0.017737 | 0.11927 | 0.11053 | 0.74323 | 1.266 | 9.9312 | 6.7245 | 7.8443 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0128.cf | 2023/09/18 10:47:23 | WS23-139-10MC-3 45-47.5cm | WS23-139-10MC-3 45-47.5cm | 0 | 0 | 892 | 4.7226 | 3004 | -25.9228 | 13C QC OK; DO NOT USE 15N | 16.533 | 0.018169 | 0.12253 | 0.1099 | 0.74113 | 1.2969 | 10.2024 | 6.7439 | 7.8669 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0153.cf | 2023/09/27 10:41:30 | WS23-139-10MC-3 47.5-50cm | WS23-139-10MC-3 47.5-50cm | 0 | 0 | 936 | 4.6415 | 3093 | -25.6954 | 13C QC OK; DO NOT USE 15N | 17.742 | 0.018747 | 0.12737 | 0.10566 | 0.71791 | 1.3381 | 10.6054 | 6.7943 | 7.9258 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0115.cf | 2023/10/03 10:29:10 | WS23-139-10MC-3 5-7.5cm | WS23-139-10MC-3 5-7.5cm | 0 | 0 | 936 | 4.6265 | 3230 | -25.8169 | 13C QC OK; DO NOT USE 15N | 16.815 | 0.018881 | 0.12502 | 0.11229 | 0.7435 | 1.3477 | 10.4096 | 6.6212 | 7.7239 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0198.cf | 2023/09/28 12:11:04 | WS23-139-10MC-3 7.5-10cm | WS23-139-10MC-3 7.5-10cm | 0 | 0 | 1037 | 4.6534 | 3494 | -25.8789 | 13C QC OK; DO NOT USE 15N | 18.667 | 0.02124 | 0.14043 | 0.11378 | 0.75229 | 1.5161 | 11.6927 | 6.6115 | 7.7125 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0160.cf | 2023/09/27 11:31:20 | WS23-139-11MC-3 0-2.5cm | WS23-139-11MC-3 0-2.5cm | 0 | 0 | 1089 | 5.0481 | 3276 | -23.1954 | 13C QC OK; DO NOT USE 15N | 16.223 | 0.021633 | 0.13351 | 0.13335 | 0.82294 | 1.5441 | 11.1162 | 6.1713 | 7.1989 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0130.cf | 2023/09/18 11:01:35 | WS23-139-11MC-3 10-12.5cm | WS23-139-11MC-3 10-12.5cm | 0 | 0 | 1030 | 5.1948 | 3110 | -23.2963 | 13C QC OK; DO NOT USE 15N | 15.717 | 0.020942 | 0.12563 | 0.13325 | 0.79934 | 1.4948 | 10.4606 | 5.999 | 6.9979 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0133.cf | 2023/09/18 11:22:56 | WS23-139-11MC-3 12.5-15cm | WS23-139-11MC-3 12.5-15cm | 0 | 0 | 1142 | 5.062 | 3472 | -23.2309 | 13C QC OK; DO NOT USE 15N | 17.732 | 0.023358 | 0.14016 | 0.13173 | 0.79041 | 1.6673 | 11.6699 | 6.0002 | 6.9994 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0123.cf | 2023/09/18 10:11:50 | WS23-139-11MC-3 15-17.5cm | WS23-139-11MC-3 15-17.5cm | 0 | 0 | 1107 | 5.0914 | 3333 | -23.3036 | 13C QC OK; DO NOT USE 15N | 17.567 | 0.022885 | 0.1358 | 0.13028 | 0.77302 | 1.6335 | 11.307 | 5.9338 | 6.9219 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0124.cf | 2023/09/18 10:18:57 | WS23-139-11MC-3 17.5-20cm | WS23-139-11MC-3 17.5-20cm | 0 | 0 | 1092 | 5.0434 | 3245 | -23.4538 | 13C QC OK; DO NOT USE 15N | 17.978 | 0.022616 | 0.13247 | 0.1258 | 0.73687 | 1.6143 | 11.0303 | 5.8575 | 6.833 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0159.cf | 2023/09/27 11:24:13 | WS23-139-11MC-3 2.5-5cm | WS23-139-11MC-3 2.5-5cm | 0 | 0 | 1072 | 5.2304 | 3232 | -23.2404 | 13C QC OK; DO NOT USE 15N | 16.03 | 0.021205 | 0.13131 | 0.13228 | 0.81915 | 1.5136 | 10.9333 | 6.1924 | 7.2236 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0161.cf | 2023/09/27 11:38:27 | WS23-139-11MC-3 20-22.5cm | WS23-139-11MC-3 20-22.5cm | 0 | 0 | 1015 | 5.1612 | 3103 | -23.5766 | 13C QC OK; DO NOT USE 15N | 16.656 | 0.020157 | 0.12591 | 0.12102 | 0.75593 | 1.4388 | 10.4835 | 6.2463 | 7.2865 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0135.cf | 2023/09/18 11:37:09 | WS23-139-11MC-3 22.5-25cm | WS23-139-11MC-3 22.5-25cm | 0 | 0 | 1100 | 4.8904 | 3433 | -23.7675 | 13C QC OK; DO NOT USE 15N | 18.011 | 0.022525 | 0.1385 | 0.12506 | 0.76899 | 1.6078 | 11.5322 | 6.1488 | 7.1728 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0155.cf | 2023/09/27 10:55:44 | WS23-139-11MC-3 25-27.5cm | WS23-139-11MC-3 25-27.5cm | 0 | 0 | 1011 | 4.9961 | 3128 | -23.7862 | 13C QC OK; DO NOT USE 15N | 16.993 | 0.020377 | 0.12853 | 0.11992 | 0.75638 | 1.4545 | 10.7021 | 6.3076 | 7.358 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0134.cf | 2023/09/18 11:30:02 | WS23-139-11MC-3 27.5-30cm | WS23-139-11MC-3 27.5-30cm | 0 | 0 | 995 | 4.9192 | 3069 | -23.8456 | 13C QC OK; DO NOT USE 15N | 16.537 | 0.020286 | 0.12404 | 0.12267 | 0.75007 | 1.448 | 10.328 | 6.1144 | 7.1326 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0132.cf | 2023/09/18 11:15:49 | WS23-139-11MC-3 30-32.5cm | WS23-139-11MC-3 30-32.5cm | 0 | 0 | 998 | 4.9762 | 3036 | -23.7428 | 13C QC OK; DO NOT USE 15N | 16.512 | 0.020317 | 0.12288 | 0.12304 | 0.74418 | 1.4502 | 10.2314 | 6.0481 | 7.0552 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0154.cf | 2023/09/27 10:48:37 | WS23-139-11MC-3 32.5-35cm | WS23-139-11MC-3 32.5-35cm | 0 | 0 | 927 | 5.064 | 2702 | -23.5805 | 13C QC OK; DO NOT USE 15N | 15.88 | 0.018345 | 0.11004 | 0.11552 | 0.69295 | 1.3094 | 9.1624 | 5.9983 | 6.9972 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0137.cf | 2023/09/18 11:51:26 | WS23-139-11MC-3 35-37.5cm | WS23-139-11MC-3 35-37.5cm | 0 | 0 | 989 | 5.0244 | 3058 | -23.8424 | 13C QC OK; DO NOT USE 15N | 17.391 | 0.020182 | 0.1231 | 0.11605 | 0.70783 | 1.4405 | 10.2496 | 6.0995 | 7.1152 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0120.cf | 2023/09/18 09:50:29 | WS23-139-11MC-3 37.5-40cm | WS23-139-11MC-3 37.5-40cm | 0 | 0 | 924 | 5.2268 | 2902 | -24.1281 | 13C QC OK; DO NOT USE 15N | 16.593 | 0.019166 | 0.12007 | 0.11551 | 0.72361 | 1.368 | 9.9974 | 6.2646 | 7.3078 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0136.cf | 2023/09/18 11:44:19 | WS23-139-11MC-3 5-7.5cm | WS23-139-11MC-3 5-7.5cm | 0 | 0 | 1054 | 5.0202 | 3197 | -23.3033 | 13C QC OK; DO NOT USE 15N | 15.957 | 0.021403 | 0.12885 | 0.13413 | 0.80747 | 1.5277 | 10.7284 | 6.0202 | 7.0227 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0156.cf | 2023/09/27 11:02:51 | WS23-139-11MC-3 7.5-10cm | WS23-139-11MC-3 7.5-10cm | 0 | 0 | 1114 | 4.653 | 3355 | -23.2504 | 13C QC OK; DO NOT USE 15N | 17.201 | 0.022369 | 0.13812 | 0.13005 | 0.80295 | 1.5967 | 11.5001 | 6.1744 | 7.2026 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0123.cf | 2023/10/03 11:26:08 | WS23-139-16MC-3 10-12.5cm | WS23-139-16MC-3 10-12.5cm | 0 | 0 | 900 | 5.025 | 2852 | -23.5544 | 13C QC OK; DO NOT USE 15N | 16.701 | 0.017805 | 0.10965 | 0.10661 | 0.65656 | 1.2709 | 9.13 | 6.1585 | 7.1841 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0119.cf | 2023/10/03 10:57:39 | WS23-139-16MC-3 12.5-15cm | WS23-139-16MC-3 12.5-15cm | 0 | 0 | 885 | 5.0606 | 2788 | -23.5386 | 13C QC OK; DO NOT USE 15N | 16.565 | 0.017552 | 0.10758 | 0.10596 | 0.64946 | 1.2528 | 8.9578 | 6.1294 | 7.1501 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0162.cf | 2023/09/27 11:45:34 | WS23-139-16MC-3 15-17.5cm | WS23-139-16MC-3 15-17.5cm | 0 | 0 | 937 | 5.2242 | 2843 | -23.5833 | 13C QC OK; DO NOT USE 15N | 17.495 | 0.018346 | 0.11556 | 0.10487 | 0.66055 | 1.3095 | 9.6222 | 6.299 | 7.3479 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0121.cf | 2023/10/03 11:11:54 | WS23-139-16MC-3 17.5-20cm | WS23-139-16MC-3 17.5-20cm | 0 | 0 | 1003 | 5.1417 | 3189 | -23.4979 | 13C QC OK; DO NOT USE 15N | 18.789 | 0.019999 | 0.12264 | 0.10644 | 0.65275 | 1.4275 | 10.2119 | 6.1324 | 7.1536 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0120.cf | 2023/10/03 11:04:46 | WS23-139-16MC-3 20-22.5cm | WS23-139-16MC-3 20-22.5cm | 0 | 0 | 943 | 5.1821 | 2991 | -23.5546 | 13C QC OK; DO NOT USE 15N | 17.78 | 0.018861 | 0.11531 | 0.10608 | 0.64855 | 1.3462 | 9.6014 | 6.1139 | 7.132 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0189.cf | 2023/09/28 11:07:02 | WS23-139-16MC-3 22.5-25cm | WS23-139-16MC-3 22.5-25cm | 0 | 0 | 887 | 5.026 | 2713 | -23.7688 | 13C QC OK; DO NOT USE 15N | 16.649 | 0.018007 | 0.10913 | 0.10816 | 0.65548 | 1.2853 | 9.0867 | 6.0605 | 7.0697 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0197.cf | 2023/09/28 12:03:57 | WS23-139-16MC-3 25-27.5cm | WS23-139-16MC-3 25-27.5cm | 0 | 0 | 920 | 5.015 | 2800 | -23.7736 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.018699 | 0.11345 | 0.10853 | 0.65847 | 1.3347 | 9.4462 | 6.0672 | 7.0776 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0124.cf | 2023/10/03 11:33:16 | WS23-139-16MC-3 27.5-30cm | WS23-139-16MC-3 27.5-30cm | 0 | 0 | 1004 | 4.9667 | 3171 | -23.4367 | 13C QC OK; DO NOT USE 15N | 18.643 | 0.02009 | 0.12199 | 0.10776 | 0.65436 | 1.434 | 10.1575 | 6.0722 | 7.0834 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0201.cf | 2023/09/28 12:32:26 | WS23-139-16MC-3 30-32.5cm | WS23-139-16MC-3 30-32.5cm | 0 | 0 | 951 | 5.3283 | 2899 | -23.7796 | 13C QC OK; DO NOT USE 15N | 17.75 | 0.01934 | 0.11716 | 0.10896 | 0.66005 | 1.3804 | 9.7551 | 6.0579 | 7.0668 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0167.cf | 2023/09/27 12:21:11 | WS23-139-16MC-3 32.5-35cm | WS23-139-16MC-3 32.5-35cm | 0 | 0 | 996 | 5.1852 | 3058 | -23.6327 | 13C QC OK; DO NOT USE 15N | 18.655 | 0.019621 | 0.12369 | 0.10518 | 0.66305 | 1.4005 | 10.2991 | 6.304 | 7.3537 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0187.cf | 2023/09/28 10:52:47 | WS23-139-16MC-3 35-36cm | WS23-139-16MC-3 35-36cm | 0 | 0 | 998 | 5.0488 | 3020 | -23.7158 | 13C QC OK; DO NOT USE 15N | 18.705 | 0.020205 | 0.12118 | 0.10802 | 0.64785 | 1.4422 | 10.09 | 5.9975 | 6.9963 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0125.cf | 2023/10/03 11:40:23 | WS23-139-17MC-3 45-47.5cm | WS23-139-17MC-3 45-47.5cm | 0 | 0 | 1550 | 4.9085 | 9124 | -26.871 | 13C QC OK; DO NOT USE 15N | 18.472 | 0.031744 | 0.36142 | 0.17185 | 1.9566 | 2.2658 | 30.093 | 11.3854 | 13.2814 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0129.cf | 2023/10/03 12:08:52 | WS23-139-17MC-3 47.5-48.75cm | WS23-139-17MC-3 47.5-48.75cm | 0 | 0 | 1294 | 5.0464 | 6054 | -26.1607 | 13C QC OK; DO NOT USE 15N | 16.697 | 0.026758 | 0.24203 | 0.16026 | 1.4496 | 1.9099 | 20.1527 | 9.0453 | 10.5515 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0116.cf | 2023/10/03 10:36:17 | WS23-139-18MC-3 0-2.5cm | WS23-139-18MC-3 0-2.5cm | 0 | 0 | 936 | 5.0125 | 3043 | -22.99 | 13C QC OK; DO NOT USE 15N | 17.304 | 0.018746 | 0.11734 | 0.10833 | 0.67809 | 1.338 | 9.7699 | 6.2593 | 7.3017 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0128.cf | 2023/10/03 12:01:44 | WS23-139-18MC-3 10-12.5cm | WS23-139-18MC-3 10-12.5cm | 0 | 0 | 824 | 4.7839 | 2611 | -23.0869 | 13C QC OK; DO NOT USE 15N | 17.429 | 0.016437 | 0.10152 | 0.094311 | 0.58248 | 1.1733 | 8.453 | 6.1762 | 7.2047 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0164.cf | 2023/09/27 11:59:49 | WS23-139-18MC-3 12.5-15cm | WS23-139-18MC-3 12.5-15cm | 0 | 0 | 1017 | 5.2125 | 3217 | -23.2717 | 13C QC OK; DO NOT USE 15N | 18.262 | 0.020279 | 0.13206 | 0.11104 | 0.72316 | 1.4475 | 10.9962 | 6.5124 | 7.5969 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0185.cf | 2023/09/28 10:38:34 | WS23-139-18MC-3 15-17.5cm | WS23-139-18MC-3 15-17.5cm | 0 | 0 | 1100 | 4.9445 | 3572 | -23.2996 | 13C QC OK; DO NOT USE 15N | 18.433 | 0.022651 | 0.14555 | 0.12288 | 0.78963 | 1.6168 | 12.1193 | 6.4258 | 7.4959 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0195.cf | 2023/09/28 11:49:42 | WS23-139-18MC-3 17.5-20cm | WS23-139-18MC-3 17.5-20cm | 0 | 0 | 915 | 5.0653 | 2842 | -23.3299 | 13C QC OK; DO NOT USE 15N | 17.305 | 0.018686 | 0.11598 | 0.10798 | 0.67018 | 1.3338 | 9.6565 | 6.2064 | 7.24 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0190.cf | 2023/09/28 11:14:09 | WS23-139-18MC-3 2.5-5cm | WS23-139-18MC-3 2.5-5cm | 0 | 0 | 1129 | 5.2451 | 3779 | -23.0664 | 13C QC OK; DO NOT USE 15N | 16.532 | 0.023288 | 0.15285 | 0.14086 | 0.92457 | 1.6622 | 12.7269 | 6.5636 | 7.6566 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0122.cf | 2023/10/03 11:19:01 | WS23-139-18MC-3 20-22.5cm | WS23-139-18MC-3 20-22.5cm | 0 | 0 | 967 | 4.8509 | 3227 | -23.0488 | 13C QC OK; DO NOT USE 15N | 18.639 | 0.019427 | 0.12544 | 0.10423 | 0.67302 | 1.3867 | 10.445 | 6.4572 | 7.5325 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0130.cf | 2023/10/03 12:15:59 | WS23-139-18MC-3 5-7.5cm | WS23-139-18MC-3 5-7.5cm | 0 | 0 | 1028 | 5.0278 | 3550 | -22.821 | 13C QC OK; DO NOT USE 15N | 16.244 | 0.020788 | 0.13736 | 0.12797 | 0.84558 | 1.4838 | 11.4368 | 6.6075 | 7.7079 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N
Acquisition-0163.cf | 2023/09/27 11:52:42 | WS23-139-18MC-3 7.5-10cm | WS23-139-18MC-3 7.5-10cm | 0 | 0 | 1013 | 5.0938 | 3323 | -23.3657 | 13C QC OK; DO NOT USE 15N | 17.172 | 0.020137 | 0.13561 | 0.11727 | 0.78969 | 1.4373 | 11.291 | 6.7342 | 7.8556 | C QC OK; N and C:N use discretion, these are acidified samples |  | N | N

## 2023-09-18a
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | Scale-Normalized Blank-corrected d15N (AT-Air) | Ampl 44 | Scale-Normalized Blank-corrected d13C (VPDB) | IRMS comment | Amount | N (mg) calculated (blank corrected) | C (mg) calculated (blank corrected) | %N calculated | %C calculated | N (µmol) | C (µmol) | C:N mass | C:N molar | EA comment | LAB USE ONLY MG Drift Correction 15N? | LAB USE ONLY MG Drift Correction 13C?
Acquisition-0112.cf | 2023/09/18 08:53:36 | BYPASS | BYPASS | 0 | 0 | 99 | 2.1733 | 31 | -30.4793 | QC OK | 0 | 0.0007345 | -8.8741e-05 | #DIV/0! | #DIV/0! | 0.052427 | -0.0073889 | -0.12082 | -0.14094 | QC OK | N | N
Acquisition-0113.cf | 2023/09/18 09:00:43 | BLANK | BLANK | 0 | 0 | 70 | -8.163 | 34 | -136.8319 | QC OK | 0 | -0.00019215 | 1.8137e-05 | #DIV/0! | #DIV/0! | -0.013715 | 0.0015101 | -0.094387 | -0.11011 | QC OK | N | N
Acquisition-0114.cf | 2023/09/18 09:07:50 | USGS 40 | USGS 40 | 0 | 0 | 739 | -4.4824 | 1568 | -26.3971 | QC OK | 0.165 | 0.015004 | 0.063908 | 9.0932 | 38.7319 | 1.0709 | 5.3212 | 4.2594 | 4.9687 | QC OK | N | N
Acquisition-0115.cf | 2023/09/18 09:14:57 | USGS 40 | USGS 40 | 0 | 0 | 3902 | -4.5575 | 8709 | -26.3828 | QC OK | 0.89 | 0.083641 | 0.35774 | 9.3979 | 40.1952 | 5.9701 | 29.7866 | 4.2771 | 4.9893 | QC OK | N | N
Acquisition-0116.cf | 2023/09/18 09:22:03 | USGS 41a | USGS 41a | 0 | 0 | 837 | 47.5764 | 1811 | 36.4984 | QC OK | 0.189 | 0.017104 | 0.073888 | 9.0497 | 39.0942 | 1.2208 | 6.1522 | 4.32 | 5.0394 | QC OK | N | N
Acquisition-0117.cf | 2023/09/18 09:29:10 | USGS 41a | USGS 41a | 0 | 0 | 4018 | 47.5235 | 9011 | 36.6015 | QC OK | 0.908 | 0.085927 | 0.36924 | 9.4634 | 40.6657 | 6.1333 | 30.7448 | 4.2972 | 5.0128 | QC OK | N | N
Acquisition-0118.cf | 2023/09/18 09:36:16 | NIST 2702 | NIST 2702 | 0 | 0 | 548 | 6.2919 | 3075 | -24.6064 | QC OK | 4.238 | 0.010589 | 0.12731 | 0.24985 | 3.0039 | 0.75581 | 10.5999 | 12.0226 | 14.0247 | QC OK | N | N
Acquisition-0119.cf | 2023/09/18 09:43:23 | NIST 2702 | NIST 2702 | 0 | 0 | 928 | 6.0654 | 5244 | -24.6601 | QC OK | 7.487 | 0.018905 | 0.22243 | 0.2525 | 2.9709 | 1.3494 | 18.5206 | 11.7659 | 13.7252 | QC OK | N | N
Acquisition-0120.cf | 2023/09/18 09:50:29 | WS23-139-11MC-3 37.5-40cm | WS23-139-11MC-3 37.5-40cm | 0 | 0 | 924 | 5.2268 | 2902 | -24.1281 | 13C QC OK; DO NOT USE 15N | 16.593 | 0.019166 | 0.12007 | 0.11551 | 0.72361 | 1.368 | 9.9974 | 6.2646 | 7.3078 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0121.cf | 2023/09/18 09:57:37 | WS23-139-10MC-3 15-17.5cm | WS23-139-10MC-3 15-17.5cm | 0 | 0 | 890 | 4.4576 | 2955 | -25.9318 | 13C QC OK; DO NOT USE 15N | 16.358 | 0.018166 | 0.12094 | 0.11105 | 0.73934 | 1.2966 | 10.07 | 6.6576 | 7.7663 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0122.cf | 2023/09/18 10:04:43 | WS23-139-10MC-3 12.5-15cm | WS23-139-10MC-3 12.5-15cm | 0 | 0 | 879 | 4.7628 | 2881 | -25.885 | 13C QC OK; DO NOT USE 15N | 16.17 | 0.017942 | 0.11795 | 0.11096 | 0.72942 | 1.2806 | 9.8207 | 6.5738 | 7.6686 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0123.cf | 2023/09/18 10:11:50 | WS23-139-11MC-3 15-17.5cm | WS23-139-11MC-3 15-17.5cm | 0 | 0 | 1107 | 5.0914 | 3333 | -23.3036 | 13C QC OK; DO NOT USE 15N | 17.567 | 0.022885 | 0.1358 | 0.13028 | 0.77302 | 1.6335 | 11.307 | 5.9338 | 6.9219 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0124.cf | 2023/09/18 10:18:57 | WS23-139-11MC-3 17.5-20cm | WS23-139-11MC-3 17.5-20cm | 0 | 0 | 1092 | 5.0434 | 3245 | -23.4538 | 13C QC OK; DO NOT USE 15N | 17.978 | 0.022616 | 0.13247 | 0.1258 | 0.73687 | 1.6143 | 11.0303 | 5.8575 | 6.833 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0125.cf | 2023/09/18 10:26:04 | WS23-139-10MC-3 30-32.5cm | WS23-139-10MC-3 30-32.5cm | 0 | 0 | 949 | 4.6771 | 3137 | -25.9889 | 13C QC OK; DO NOT USE 15N | 17.341 | 0.019489 | 0.12834 | 0.11239 | 0.74009 | 1.3911 | 10.686 | 6.5851 | 7.6817 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0126.cf | 2023/09/18 10:33:11 | WS23-139-10MC-3 40-42.5cm | WS23-139-10MC-3 40-42.5cm | 0 | 0 | 926 | 4.8872 | 3117 | -26.0041 | 13C QC OK; DO NOT USE 15N | 16.89 | 0.018799 | 0.12669 | 0.1113 | 0.75008 | 1.3418 | 10.5487 | 6.7391 | 7.8614 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0127.cf | 2023/09/18 10:40:17 | WS23-139-10MC-3 27.5-30cm | WS23-139-10MC-3 27.5-30cm | 0 | 0 | 921 | 4.7354 | 3064 | -25.9432 | 13C QC OK; DO NOT USE 15N | 16.929 | 0.018884 | 0.1251 | 0.11155 | 0.73896 | 1.3479 | 10.4161 | 6.6244 | 7.7276 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0128.cf | 2023/09/18 10:47:23 | WS23-139-10MC-3 45-47.5cm | WS23-139-10MC-3 45-47.5cm | 0 | 0 | 892 | 4.7226 | 3004 | -25.9228 | 13C QC OK; DO NOT USE 15N | 16.533 | 0.018169 | 0.12253 | 0.1099 | 0.74113 | 1.2969 | 10.2024 | 6.7439 | 7.8669 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0129.cf | 2023/09/18 10:54:29 | WS23-139-10MC-3 42.5-45cm | WS23-139-10MC-3 42.5-45cm | 0 | 0 | 873 | 4.8835 | 2933 | -26.0058 | 13C QC OK; DO NOT USE 15N | 16.048 | 0.017737 | 0.11927 | 0.11053 | 0.74323 | 1.266 | 9.9312 | 6.7245 | 7.8443 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0130.cf | 2023/09/18 11:01:35 | WS23-139-11MC-3 10-12.5cm | WS23-139-11MC-3 10-12.5cm | 0 | 0 | 1030 | 5.1948 | 3110 | -23.2963 | 13C QC OK; DO NOT USE 15N | 15.717 | 0.020942 | 0.12563 | 0.13325 | 0.79934 | 1.4948 | 10.4606 | 5.999 | 6.9979 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0131.cf | 2023/09/18 11:08:42 | WS23-139-10MC-3 35-37.5cm | WS23-139-10MC-3 35-37.5cm | 0 | 0 | 884 | 4.8156 | 2994 | -26.0475 | 13C QC OK; DO NOT USE 15N | 15.915 | 0.017863 | 0.12163 | 0.11224 | 0.76425 | 1.275 | 10.1275 | 6.809 | 7.9429 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0132.cf | 2023/09/18 11:15:49 | WS23-139-11MC-3 30-32.5cm | WS23-139-11MC-3 30-32.5cm | 0 | 0 | 998 | 4.9762 | 3036 | -23.7428 | 13C QC OK; DO NOT USE 15N | 16.512 | 0.020317 | 0.12288 | 0.12304 | 0.74418 | 1.4502 | 10.2314 | 6.0481 | 7.0552 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0133.cf | 2023/09/18 11:22:56 | WS23-139-11MC-3 12.5-15cm | WS23-139-11MC-3 12.5-15cm | 0 | 0 | 1142 | 5.062 | 3472 | -23.2309 | 13C QC OK; DO NOT USE 15N | 17.732 | 0.023358 | 0.14016 | 0.13173 | 0.79041 | 1.6673 | 11.6699 | 6.0002 | 6.9994 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0134.cf | 2023/09/18 11:30:02 | WS23-139-11MC-3 27.5-30cm | WS23-139-11MC-3 27.5-30cm | 0 | 0 | 995 | 4.9192 | 3069 | -23.8456 | 13C QC OK; DO NOT USE 15N | 16.537 | 0.020286 | 0.12404 | 0.12267 | 0.75007 | 1.448 | 10.328 | 6.1144 | 7.1326 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0135.cf | 2023/09/18 11:37:09 | WS23-139-11MC-3 22.5-25cm | WS23-139-11MC-3 22.5-25cm | 0 | 0 | 1100 | 4.8904 | 3433 | -23.7675 | 13C QC OK; DO NOT USE 15N | 18.011 | 0.022525 | 0.1385 | 0.12506 | 0.76899 | 1.6078 | 11.5322 | 6.1488 | 7.1728 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0136.cf | 2023/09/18 11:44:19 | WS23-139-11MC-3 5-7.5cm | WS23-139-11MC-3 5-7.5cm | 0 | 0 | 1054 | 5.0202 | 3197 | -23.3033 | 13C QC OK; DO NOT USE 15N | 15.957 | 0.021403 | 0.12885 | 0.13413 | 0.80747 | 1.5277 | 10.7284 | 6.0202 | 7.0227 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0137.cf | 2023/09/18 11:51:26 | WS23-139-11MC-3 35-37.5cm | WS23-139-11MC-3 35-37.5cm | 0 | 0 | 989 | 5.0244 | 3058 | -23.8424 | 13C QC OK; DO NOT USE 15N | 17.391 | 0.020182 | 0.1231 | 0.11605 | 0.70783 | 1.4405 | 10.2496 | 6.0995 | 7.1152 | C QC OK; DO NOT USE N, C:N | N | N
Acquisition-0138.cf | 2023/09/18 11:58:32 | NIST 2702 | NIST 2702 | 0 | 0 | 707 | 6.0652 | 4277 | -24.6959 | QC OK | 5.46 | 0.013945 | 0.17524 | 0.25541 | 3.2096 | 0.99538 | 14.5914 | 12.5665 | 14.6591 | QC OK | N | N
Acquisition-0139.cf | 2023/09/18 12:05:38 | NIST 2702 | NIST 2702 | 0 | 0 | 1227 | 6.0247 | 7433 | -24.6886 | QC OK | 9.838 | 0.02518 | 0.30653 | 0.25595 | 3.1158 | 1.7973 | 25.5232 | 12.1737 | 14.201 | QC OK | N | N
Acquisition-0140.cf | 2023/09/18 12:12:44 | NIST 2702 | NIST 2702 | 0 | 0 | 949 | 6.1039 | 5769 | -24.5965 | QC OK | 7.535 | 0.019182 | 0.23703 | 0.25458 | 3.1457 | 1.3692 | 19.7359 | 12.3566 | 14.4143 | QC OK | N | N
Acquisition-0141.cf | 2023/09/18 12:19:51 | USGS 40 | USGS 40 | 0 | 0 | 2895 | -4.4914 | 6662 | -26.3303 | QC OK | 0.661 | 0.061799 | 0.27292 | 9.3493 | 41.2882 | 4.4111 | 22.724 | 4.4162 | 5.1516 | QC OK | N | N
Acquisition-0142.cf | 2023/09/18 12:26:57 | USGS 41a | USGS 41a | 0 | 0 | 3213 | 47.5266 | 7315 | 36.558 | QC OK | 0.727 | 0.068681 | 0.30002 | 9.4472 | 41.2679 | 4.9023 | 24.9806 | 4.3682 | 5.0957 | QC OK | N | N
Acquisition-0143.cf | 2023/09/18 12:34:03 | BLANK | BLANK | 0 | 0 | 59 | 1.9035 | 35 | 12.6316 | QC OK | 0 | -0.00054235 | 7.0604e-05 | #DIV/0! | #DIV/0! | -0.038711 | 0.0058788 | -0.13018 | -0.15186 | QC OK | N | N

## 2023-09-27a
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | Scale-Normalized Blank-corrected d15N (AT-Air) | Ampl 44 | Scale-Normalized Blank-corrected d13C (VPDB) | IRMS comment | Amount | N (mg) calculated (blank corrected) | C (mg) calculated (blank corrected) | %N calculated | %C calculated | N (µmol) | C (µmol) | C:N mass | C:N molar | EA comment | LAB USE ONLY MG Drift Correction 15N? | LAB USE ONLY MG Drift Correction 13C?
Acquisition-0144.cf | 2023/09/27 09:37:27 | BYPASS | BYPASS | 0 | 0 | 125 | 1.2946 | 31 | -20.9756 | QC OK | 0 | 0.0012094 | -0.00013937 | #DIV/0! | #DIV/0! | 0.086325 | -0.011605 | -0.11524 | -0.13443 | QC OK | N | N
Acquisition-0145.cf | 2023/09/27 09:44:33 | BLANK | BLANK | 0 | 0 | 71 | -9.1977 | 30 | -19.3284 | QC OK | 0 | -0.00044701 | -0.00019383 | #DIV/0! | #DIV/0! | -0.031906 | -0.016139 | 0.43361 | 0.50582 | QC OK | N | N
Acquisition-0146.cf | 2023/09/27 09:51:39 | USGS 40 | USGS 40 | 0 | 0 | 982 | -4.5745 | 2154 | -26.3754 | QC OK | 0.221 | 0.019758 | 0.088267 | 8.9402 | 39.9396 | 1.4103 | 7.3494 | 4.4674 | 5.2114 | QC OK | N | N
Acquisition-0147.cf | 2023/09/27 09:58:47 | USGS 40 | USGS 40 | 0 | 0 | 3750 | -4.463 | 8188 | -26.4039 | QC OK | 0.844 | 0.078958 | 0.33737 | 9.3552 | 39.9732 | 5.6359 | 28.0911 | 4.2728 | 4.9844 | QC OK | N | N
Acquisition-0148.cf | 2023/09/27 10:05:54 | USGS 41a | USGS 41a | 0 | 0 | 3799 | 47.3014 | 8509 | 36.6963 | QC OK | 0.85 | 0.07989 | 0.3503 | 9.3988 | 41.2123 | 5.7024 | 29.1677 | 4.3848 | 5.115 | QC OK | N | N
Acquisition-0149.cf | 2023/09/27 10:13:00 | USGS 41a | USGS 41a | 0 | 0 | 694 | 47.7961 | 1493 | 36.403 | QC OK | 0.153 | 0.013416 | 0.060787 | 8.7687 | 39.7303 | 0.95761 | 5.0614 | 4.5309 | 5.2854 | QC OK | N | N
Acquisition-0150.cf | 2023/09/27 10:20:09 | NIST 2702 | NIST 2702 | 0 | 0 | 526 | 6.0101 | 2963 | -24.4605 | QC OK | 3.991 | 0.0096145 | 0.12255 | 0.24091 | 3.0706 | 0.68626 | 10.2039 | 12.7461 | 14.8687 | QC OK | N | N
Acquisition-0151.cf | 2023/09/27 10:27:16 | NIST 2702 | NIST 2702 | 0 | 0 | 934 | 6.2467 | 5321 | -24.6504 | QC OK | 7.401 | 0.018387 | 0.22323 | 0.24843 | 3.0163 | 1.3124 | 18.5873 | 12.1411 | 14.163 | QC OK | N | N
Acquisition-0152.cf | 2023/09/27 10:34:22 | WS23-139-10MC-3 37.5-40cm | WS23-139-10MC-3 37.5-40cm | 0 | 0 | 921 | 4.4286 | 3155 | -26.0594 | 13C QC OK; DO NOT USE 15N | 16.74 | 0.01829 | 0.12956 | 0.10926 | 0.77395 | 1.3055 | 10.7876 | 7.0835 | 8.2631 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0153.cf | 2023/09/27 10:41:30 | WS23-139-10MC-3 47.5-50cm | WS23-139-10MC-3 47.5-50cm | 0 | 0 | 936 | 4.6415 | 3093 | -25.6954 | 13C QC OK; DO NOT USE 15N | 17.742 | 0.018747 | 0.12737 | 0.10566 | 0.71791 | 1.3381 | 10.6054 | 6.7943 | 7.9258 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0154.cf | 2023/09/27 10:48:37 | WS23-139-11MC-3 32.5-35cm | WS23-139-11MC-3 32.5-35cm | 0 | 0 | 927 | 5.064 | 2702 | -23.5805 | 13C QC OK; DO NOT USE 15N | 15.88 | 0.018345 | 0.11004 | 0.11552 | 0.69295 | 1.3094 | 9.1624 | 5.9983 | 6.9972 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0155.cf | 2023/09/27 10:55:44 | WS23-139-11MC-3 25-27.5cm | WS23-139-11MC-3 25-27.5cm | 0 | 0 | 1011 | 4.9961 | 3128 | -23.7862 | 13C QC OK; DO NOT USE 15N | 16.993 | 0.020377 | 0.12853 | 0.11992 | 0.75638 | 1.4545 | 10.7021 | 6.3076 | 7.358 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0156.cf | 2023/09/27 11:02:51 | WS23-139-11MC-3 7.5-10cm | WS23-139-11MC-3 7.5-10cm | 0 | 0 | 1114 | 4.653 | 3355 | -23.2504 | 13C QC OK; DO NOT USE 15N | 17.201 | 0.022369 | 0.13812 | 0.13005 | 0.80295 | 1.5967 | 11.5001 | 6.1744 | 7.2026 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0157.cf | 2023/09/27 11:09:59 | WS23-139-10MC-3 32.5-35cm | WS23-139-10MC-3 32.5-35cm | 0 | 0 | 967 | 4.5913 | 3262 | -25.9362 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.019154 | 0.1325 | 0.11117 | 0.76906 | 1.3672 | 11.0326 | 6.9178 | 8.0698 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0158.cf | 2023/09/27 11:17:06 | WS23-139-10MC-3 10-12.5cm | WS23-139-10MC-3 10-12.5cm | 0 | 0 | 992 | 4.7649 | 3484 | -25.9944 | 13C QC OK; DO NOT USE 15N | 17.631 | 0.019684 | 0.14166 | 0.11164 | 0.80349 | 1.405 | 11.7954 | 7.1969 | 8.3954 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0159.cf | 2023/09/27 11:24:13 | WS23-139-11MC-3 2.5-5cm | WS23-139-11MC-3 2.5-5cm | 0 | 0 | 1072 | 5.2304 | 3232 | -23.2404 | 13C QC OK; DO NOT USE 15N | 16.03 | 0.021205 | 0.13131 | 0.13228 | 0.81915 | 1.5136 | 10.9333 | 6.1924 | 7.2236 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0160.cf | 2023/09/27 11:31:20 | WS23-139-11MC-3 0-2.5cm | WS23-139-11MC-3 0-2.5cm | 0 | 0 | 1089 | 5.0481 | 3276 | -23.1954 | 13C QC OK; DO NOT USE 15N | 16.223 | 0.021633 | 0.13351 | 0.13335 | 0.82294 | 1.5441 | 11.1162 | 6.1713 | 7.1989 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0161.cf | 2023/09/27 11:38:27 | WS23-139-11MC-3 20-22.5cm | WS23-139-11MC-3 20-22.5cm | 0 | 0 | 1015 | 5.1612 | 3103 | -23.5766 | 13C QC OK; DO NOT USE 15N | 16.656 | 0.020157 | 0.12591 | 0.12102 | 0.75593 | 1.4388 | 10.4835 | 6.2463 | 7.2865 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0162.cf | 2023/09/27 11:45:34 | WS23-139-16MC-3 15-17.5cm | WS23-139-16MC-3 15-17.5cm | 0 | 0 | 937 | 5.2242 | 2843 | -23.5833 | 13C QC OK; DO NOT USE 15N | 17.495 | 0.018346 | 0.11556 | 0.10487 | 0.66055 | 1.3095 | 9.6222 | 6.299 | 7.3479 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0163.cf | 2023/09/27 11:52:42 | WS23-139-18MC-3 7.5-10cm | WS23-139-18MC-3 7.5-10cm | 0 | 0 | 1013 | 5.0938 | 3323 | -23.3657 | 13C QC OK; DO NOT USE 15N | 17.172 | 0.020137 | 0.13561 | 0.11727 | 0.78969 | 1.4373 | 11.291 | 6.7342 | 7.8556 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0164.cf | 2023/09/27 11:59:49 | WS23-139-18MC-3 12.5-15cm | WS23-139-18MC-3 12.5-15cm | 0 | 0 | 1017 | 5.2125 | 3217 | -23.2717 | 13C QC OK; DO NOT USE 15N | 18.262 | 0.020279 | 0.13206 | 0.11104 | 0.72316 | 1.4475 | 10.9962 | 6.5124 | 7.5969 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0165.cf | 2023/09/27 12:06:56 | WS23-139-01MC-3 5-7.5cm | WS23-139-01MC-3 5-7.5cm | 0 | 0 | 364 | 4.3964 | 1778 | -27.7181 | 13C QC OK; DO NOT USE 15N | 16.748 | 0.0060456 | 0.072292 | 0.036097 | 0.43165 | 0.43152 | 6.0194 | 11.9579 | 13.9492 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0166.cf | 2023/09/27 12:14:04 | WS23-139-01MC-3 22.5-25cm | WS23-139-01MC-3 22.5-25cm | 0 | 0 | 302 | 4.0424 | 684 | -25.7423 | 13C QC OK; DO NOT USE 15N | 17.698 | 0.0046911 | 0.02699 | 0.026507 | 0.1525 | 0.33484 | 2.2473 | 5.7533 | 6.7114 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0167.cf | 2023/09/27 12:21:11 | WS23-139-16MC-3 32.5-35cm | WS23-139-16MC-3 32.5-35cm | 0 | 0 | 996 | 5.1852 | 3058 | -23.6327 | 13C QC OK; DO NOT USE 15N | 18.655 | 0.019621 | 0.12369 | 0.10518 | 0.66305 | 1.4005 | 10.2991 | 6.304 | 7.3537 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0168.cf | 2023/09/27 12:28:18 | WS23-139-01MC-3 12.5-15cm | WS23-139-01MC-3 12.5-15cm | 0 | 0 | 296 | 3.6666 | 588 | -25.6388 | 13C QC OK; DO NOT USE 15N | 17.957 | 0.0045648 | 0.023059 | 0.02542 | 0.12841 | 0.32582 | 1.92 | 5.0516 | 5.8929 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0169.cf | 2023/09/27 12:35:26 | WS23-139-10MC-3 2.5-5.0cm | WS23-139-10MC-3 2.5-5.0cm | 0 | 0 | 1084 | 4.7995 | 3819 | -25.836 | 13C QC OK; DO NOT USE 15N | 18.758 | 0.021591 | 0.15526 | 0.1151 | 0.8277 | 1.5411 | 12.9276 | 7.191 | 8.3885 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0170.cf | 2023/09/27 12:42:33 | NIST 2702 | NIST 2702 | 0 | 0 | 561 | 5.9061 | 3268 | -24.7404 | QC OK | 4.28 | 0.01038 | 0.13397 | 0.24251 | 3.1301 | 0.74087 | 11.1547 | 12.9069 | 15.0562 | QC OK | N | N
Acquisition-0171.cf | 2023/09/27 12:49:40 | NIST 2702 | NIST 2702 | 0 | 0 | 861 | 5.9738 | 5226 | -24.5519 | QC OK | 6.791 | 0.016745 | 0.21585 | 0.24657 | 3.1785 | 1.1952 | 17.9725 | 12.8906 | 15.0373 | QC OK | N | N
Acquisition-0172.cf | 2023/09/27 12:56:48 | NIST 2702 | NIST 2702 | 0 | 0 | 780 | 6.0266 | 4530 | -24.7054 | QC OK | 6.13 | 0.015096 | 0.18771 | 0.24627 | 3.0621 | 1.0775 | 15.6294 | 12.4343 | 14.505 | QC OK | N | N
Acquisition-0173.cf | 2023/09/27 13:03:55 | USGS 40 | USGS 40 | 0 | 0 | 1958 | -4.6051 | 4600 | -26.2349 | QC OK | 0.447 | 0.040848 | 0.1912 | 9.1383 | 42.7744 | 2.9157 | 15.9202 | 4.6808 | 5.4603 | QC OK | N | N
Acquisition-0174.cf | 2023/09/27 13:11:01 | USGS 41a | USGS 41a | 0 | 0 | 1781 | 47.2528 | 4172 | 34.9953 | QC OK | 0.404 | 0.036908 | 0.17177 | 9.1356 | 42.5168 | 2.6344 | 14.3021 | 4.6539 | 5.429 | QC OK | N | N
Acquisition-0175.cf | 2023/09/27 13:18:08 | BLANK | BLANK | 0 | 0 | 60 | 2.7667 | 43 | -16.7836 | QC OK | 0 | -0.00076241 | 0.0003332 | #DIV/0! | #DIV/0! | -0.054419 | 0.027743 | -0.43703 | -0.50981 | QC OK | N | N

## 2023-09-28a
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | Scale-Normalized Blank-corrected d15N (AT-Air) | Ampl 44 | Scale-Normalized Blank-corrected d13C (VPDB) | IRMS comment | Amount | N (mg) calculated (blank corrected) | C (mg) calculated (blank corrected) | %N calculated | %C calculated | N (µmol) | C (µmol) | C:N mass | C:N molar | EA comment | LAB USE ONLY MG Drift Correction 15N? | LAB USE ONLY MG Drift Correction 13C?
Acquisition-0176.cf | 2023/09/28 09:34:31 | BYPASS | BYPASS | 0 | 0 | 71 | 15.4379 | 31 | -35.6875 | QC OK | 0 | 0.00015791 | -5.9928e-05 | #DIV/0! | #DIV/0! | 0.011271 | -0.0049899 | -0.37951 | -0.44271 | QC OK | N | N
Acquisition-0177.cf | 2023/09/28 09:41:39 | BLANK | BLANK | 0 | 0 | 65 | -80.3032 | 31 | -25.5363 | QC OK | 0 | -1.9458e-05 | -6.1862e-05 | #DIV/0! | #DIV/0! | -0.0013889 | -0.0051508 | 3.1792 | 3.7086 | QC OK | N | N
Acquisition-0178.cf | 2023/09/28 09:48:46 | USGS 40 | USGS 40 | 0 | 0 | 4117 | -4.5925 | 9079 | -26.3803 | QC OK | 0.926 | 0.087266 | 0.37212 | 9.424 | 40.1859 | 6.2289 | 30.9843 | 4.2642 | 4.9743 | QC OK | N | N
Acquisition-0179.cf | 2023/09/28 09:55:52 | USGS 40 | USGS 40 | 0 | 0 | 818 | -4.4464 | 1766 | -26.3995 | QC OK | 0.181 | 0.016763 | 0.071439 | 9.2615 | 39.4688 | 1.1965 | 5.9483 | 4.2616 | 4.9713 | QC OK | N | N
Acquisition-0180.cf | 2023/09/28 10:02:59 | USGS 41a | USGS 41a | 0 | 0 | 804 | 47.3944 | 1729 | 36.4603 | QC OK | 0.177 | 0.016437 | 0.069759 | 9.2862 | 39.4117 | 1.1732 | 5.8084 | 4.2441 | 4.9509 | QC OK | N | N
Acquisition-0181.cf | 2023/09/28 10:10:06 | USGS 41a | USGS 41a | 0 | 0 | 4056 | 47.7045 | 9037 | 36.6395 | QC OK | 0.899 | 0.085626 | 0.36801 | 9.5246 | 40.9358 | 6.1118 | 30.6422 | 4.2979 | 5.0136 | QC OK | N | N
Acquisition-0182.cf | 2023/09/28 10:17:13 | NIST 2702 | NIST 2702 | 0 | 0 | 450 | 6.0005 | 2497 | -24.6655 | QC OK | 3.359 | 0.0085672 | 0.10214 | 0.25505 | 3.0409 | 0.61151 | 8.5049 | 11.9225 | 13.908 | QC OK | N | N
Acquisition-0183.cf | 2023/09/28 10:24:20 | NIST 2702 | NIST 2702 | 0 | 0 | 842 | 5.9908 | 4739 | -24.564 | QC OK | 6.616 | 0.016992 | 0.19983 | 0.25684 | 3.0204 | 1.2129 | 16.6388 | 11.7602 | 13.7186 | QC OK | N | N
Acquisition-0184.cf | 2023/09/28 10:31:28 | WS23-139-10MC-3 22.5-25cm | WS23-139-10MC-3 22.5-25cm | 0 | 0 | 1014 | 5.099 | 3390 | -25.9015 | 13C QC OK; DO NOT USE 15N | 18.597 | 0.021035 | 0.13884 | 0.11311 | 0.74658 | 1.5014 | 11.5604 | 6.6005 | 7.6997 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0185.cf | 2023/09/28 10:38:34 | WS23-139-18MC-3 15-17.5cm | WS23-139-18MC-3 15-17.5cm | 0 | 0 | 1100 | 4.9445 | 3572 | -23.2996 | 13C QC OK; DO NOT USE 15N | 18.433 | 0.022651 | 0.14555 | 0.12288 | 0.78963 | 1.6168 | 12.1193 | 6.4258 | 7.4959 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0186.cf | 2023/09/28 10:45:42 | WS23-139-01MC-3 27.5-30cm | WS23-139-01MC-3 27.5-30cm | 0 | 0 | 300 | 4.3353 | 712 | -25.9516 | 13C QC OK; DO NOT USE 15N | 17.828 | 0.0053308 | 0.028472 | 0.029901 | 0.1597 | 0.3805 | 2.3707 | 5.341 | 6.2304 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0187.cf | 2023/09/28 10:52:47 | WS23-139-16MC-3 35-36cm | WS23-139-16MC-3 35-36cm | 0 | 0 | 998 | 5.0488 | 3020 | -23.7158 | 13C QC OK; DO NOT USE 15N | 18.705 | 0.020205 | 0.12118 | 0.10802 | 0.64785 | 1.4422 | 10.09 | 5.9975 | 6.9963 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0188.cf | 2023/09/28 10:59:55 | WS23-139-01MC-3 0-2.5cm | WS23-139-01MC-3 0-2.5cm | 0 | 0 | 385 | 4.0457 | 1692 | -27.2857 | 13C QC OK; DO NOT USE 15N | 17.14 | 0.007201 | 0.068228 | 0.042013 | 0.39806 | 0.51399 | 5.6809 | 9.4747 | 11.0525 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0189.cf | 2023/09/28 11:07:02 | WS23-139-16MC-3 22.5-25cm | WS23-139-16MC-3 22.5-25cm | 0 | 0 | 887 | 5.026 | 2713 | -23.7688 | 13C QC OK; DO NOT USE 15N | 16.649 | 0.018007 | 0.10913 | 0.10816 | 0.65548 | 1.2853 | 9.0867 | 6.0605 | 7.0697 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0190.cf | 2023/09/28 11:14:09 | WS23-139-18MC-3 2.5-5cm | WS23-139-18MC-3 2.5-5cm | 0 | 0 | 1129 | 5.2451 | 3779 | -23.0664 | 13C QC OK; DO NOT USE 15N | 16.532 | 0.023288 | 0.15285 | 0.14086 | 0.92457 | 1.6622 | 12.7269 | 6.5636 | 7.6566 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0191.cf | 2023/09/28 11:21:17 | WS23-139-10MC-3 20-22.5cm | WS23-139-10MC-3 20-22.5cm | 0 | 0 | 984 | 4.6768 | 3256 | -25.8781 | 13C QC OK; DO NOT USE 15N | 17.743 | 0.020275 | 0.13143 | 0.11427 | 0.74075 | 1.4472 | 10.9435 | 6.4825 | 7.562 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0192.cf | 2023/09/28 11:28:24 | WS23-139-01MC-3 2.5-5cm | WS23-139-01MC-3 2.5-5cm | 0 | 0 | 226 | 4.2666 | 646 | -26.8383 | 13C QC OK; DO NOT USE 15N | 16.612 | 0.0036941 | 0.025452 | 0.022237 | 0.15322 | 0.26367 | 2.1193 | 6.89 | 8.0374 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0193.cf | 2023/09/28 11:35:29 | WS23-139-01MC-3 25-27.5cm | WS23-139-01MC-3 25-27.5cm | 0 | 0 | 348 | 3.9236 | 815 | -26.0453 | 13C QC OK; DO NOT USE 15N | 18.559 | 0.0063625 | 0.032427 | 0.034282 | 0.17472 | 0.45414 | 2.7 | 5.0966 | 5.9454 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0194.cf | 2023/09/28 11:42:35 | WS23-139-10MC-3 17.5-20cm | WS23-139-10MC-3 17.5-20cm | 0 | 0 | 993 | 4.6877 | 3310 | -25.834 | 13C QC OK; DO NOT USE 15N | 17.982 | 0.020398 | 0.13342 | 0.11344 | 0.74198 | 1.456 | 11.1092 | 6.5409 | 7.6301 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0195.cf | 2023/09/28 11:49:42 | WS23-139-18MC-3 17.5-20cm | WS23-139-18MC-3 17.5-20cm | 0 | 0 | 915 | 5.0653 | 2842 | -23.3299 | 13C QC OK; DO NOT USE 15N | 17.305 | 0.018686 | 0.11598 | 0.10798 | 0.67018 | 1.3338 | 9.6565 | 6.2064 | 7.24 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0196.cf | 2023/09/28 11:56:49 | WS23-139-10MC-3 0-2.5cm | WS23-139-10MC-3 0-2.5cm | 0 | 0 | 1046 | 4.7378 | 3455 | -25.5252 | 13C QC OK; DO NOT USE 15N | 17.36 | 0.021343 | 0.13896 | 0.12295 | 0.80045 | 1.5234 | 11.5702 | 6.5106 | 7.5948 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0197.cf | 2023/09/28 12:03:57 | WS23-139-16MC-3 25-27.5cm | WS23-139-16MC-3 25-27.5cm | 0 | 0 | 920 | 5.015 | 2800 | -23.7736 | 13C QC OK; DO NOT USE 15N | 17.229 | 0.018699 | 0.11345 | 0.10853 | 0.65847 | 1.3347 | 9.4462 | 6.0672 | 7.0776 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0198.cf | 2023/09/28 12:11:04 | WS23-139-10MC-3 7.5-10cm | WS23-139-10MC-3 7.5-10cm | 0 | 0 | 1037 | 4.6534 | 3494 | -25.8789 | 13C QC OK; DO NOT USE 15N | 18.667 | 0.02124 | 0.14043 | 0.11378 | 0.75229 | 1.5161 | 11.6927 | 6.6115 | 7.7125 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0199.cf | 2023/09/28 12:18:11 | WS23-139-01MC-3 10-12.5cm | WS23-139-01MC-3 10-12.5cm | 0 | 0 | 405 | 4.4545 | 1279 | -26.5615 | 13C QC OK; DO NOT USE 15N | 18.039 | 0.0075749 | 0.051233 | 0.041992 | 0.28401 | 0.54068 | 4.2659 | 6.7636 | 7.8899 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0200.cf | 2023/09/28 12:25:18 | WS23-139-10MC-3 25-27.5cm | WS23-139-10MC-3 25-27.5cm | 0 | 0 | 957 | 4.6191 | 3173 | -25.9636 | 13C QC OK; DO NOT USE 15N | 17.002 | 0.01948 | 0.12782 | 0.11457 | 0.75181 | 1.3904 | 10.6431 | 6.5618 | 7.6545 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0201.cf | 2023/09/28 12:32:26 | WS23-139-16MC-3 30-32.5cm | WS23-139-16MC-3 30-32.5cm | 0 | 0 | 951 | 5.3283 | 2899 | -23.7796 | 13C QC OK; DO NOT USE 15N | 17.75 | 0.01934 | 0.11716 | 0.10896 | 0.66005 | 1.3804 | 9.7551 | 6.0579 | 7.0668 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0202.cf | 2023/09/28 12:39:33 | NIST 2702 | NIST 2702 | 0 | 0 | 1233 | 6.069 | 7241 | -24.7246 | QC OK | 9.791 | 0.025237 | 0.29844 | 0.25775 | 3.0481 | 1.8013 | 24.8491 | 11.8256 | 13.7949 | QC OK | N | N
Acquisition-0203.cf | 2023/09/28 12:46:40 | NIST 2702 | NIST 2702 | 0 | 0 | 791 | 5.9469 | 4967 | -24.8198 | QC OK | 6.117 | 0.015872 | 0.20422 | 0.25947 | 3.3386 | 1.1329 | 17.0046 | 12.867 | 15.0097 | QC OK | N | N
Acquisition-0204.cf | 2023/09/28 12:53:46 | NIST 2702 | NIST 2702 | 0 | 0 | 646 | 5.948 | 3872 | -24.8625 | QC OK | 4.978 | 0.012774 | 0.1579 | 0.2566 | 3.172 | 0.91175 | 13.1475 | 12.3616 | 14.4201 | QC OK | N | N
Acquisition-0205.cf | 2023/09/28 13:00:53 | USGS 40 | USGS 40 | 0 | 0 | 1261 | -4.5882 | 2874 | -26.4835 | QC OK | 0.285 | 0.026395 | 0.11689 | 9.2614 | 41.0125 | 1.884 | 9.7324 | 4.4283 | 5.1657 | QC OK | N | N
Acquisition-0206.cf | 2023/09/28 13:08:01 | USGS 41a | USGS 41a | 0 | 0 | 1633 | 47.4628 | 3724 | 36.2354 | QC OK | 0.368 | 0.034446 | 0.15153 | 9.3604 | 41.1772 | 2.4587 | 12.6172 | 4.3991 | 5.1316 | QC OK | N | N
Acquisition-0207.cf | 2023/09/28 13:15:08 | BLANK | BLANK | 0 | 0 | 61 | 24.8956 | 36 | -30.955 | QC OK | 0 | -0.00013845 | 0.00012179 | #DIV/0! | #DIV/0! | -0.0098824 | 0.010141 | -0.87965 | -1.0261 | QC OK | N | N

## 2023-10-03a
FileHeader: Filename | Time Code | Identifier 1 | Identifier 2 | adjusted ID 1 | adjusted ID 2 | Ampl 28 | Scale-Normalized Blank-corrected d15N (AT-Air) | Ampl 44 | Scale-Normalized Blank-corrected d13C (VPDB) | IRMS comment | Amount | N (mg) calculated (blank corrected) | C (mg) calculated (blank corrected) | %N calculated | %C calculated | N (µmol) | C (µmol) | C:N mass | C:N molar | EA comment | LAB USE ONLY MG Drift Correction 15N? | LAB USE ONLY MG Drift Correction 13C?
Acquisition-0107.cf | 2023/10/03 09:32:16 | BYPASS | BYPASS | 0 | 0 | 90 | 8.5647 | 32 | -3.7993 | QC OK | 0 | 0.00042093 | -9.2403e-05 | #DIV/0! | #DIV/0! | 0.030045 | -0.0076938 | -0.21952 | -0.25607 | QC OK | N | N
Acquisition-0108.cf | 2023/10/03 09:39:22 | BLANK | BLANK | 0 | 0 | 68 | -7.0595 | 30 | -25.6286 | QC OK | 0 | -0.00023977 | -0.00014143 | #DIV/0! | #DIV/0! | -0.017114 | -0.011776 | 0.58987 | 0.68811 | QC OK | N | N
Acquisition-0109.cf | 2023/10/03 09:46:28 | USGS 40 | USGS 40 | 0 | 0 | 789 | -4.3489 | 1759 | -26.3949 | QC OK | 0.175 | 0.015842 | 0.067959 | 9.0527 | 38.8339 | 1.1308 | 5.6586 | 4.2898 | 5.0041 | QC OK | N | N
Acquisition-0110.cf | 2023/10/03 09:53:35 | USGS 40 | USGS 40 | 0 | 0 | 4266 | -4.6897 | 9880 | -26.3848 | QC OK | 0.956 | 0.089697 | 0.38493 | 9.3825 | 40.2645 | 6.4024 | 32.0507 | 4.2914 | 5.0061 | QC OK | N | N
Acquisition-0111.cf | 2023/10/03 10:00:42 | USGS 41a | USGS 41a | 0 | 0 | 808 | 47.6369 | 1813 | 36.4602 | QC OK | 0.176 | 0.016129 | 0.069811 | 9.1645 | 39.6654 | 1.1513 | 5.8127 | 4.3282 | 5.0489 | QC OK | N | N
Acquisition-0112.cf | 2023/10/03 10:07:49 | USGS 41a | USGS 41a | 0 | 0 | 4317 | 47.4616 | 9990 | 36.6395 | QC OK | 0.952 | 0.090507 | 0.38828 | 9.507 | 40.7854 | 6.4602 | 32.3295 | 4.29 | 5.0044 | QC OK | N | N
Acquisition-0113.cf | 2023/10/03 10:14:55 | NIST 2702 | NIST 2702 | 0 | 0 | 835 | 6.0868 | 5003 | -24.4885 | QC OK | 6.475 | 0.016442 | 0.19824 | 0.25393 | 3.0617 | 1.1736 | 16.5065 | 12.0571 | 14.0649 | QC OK | N | N
Acquisition-0114.cf | 2023/10/03 10:22:02 | NIST 2702 | NIST 2702 | 0 | 0 | 539 | 6.15 | 3161 | -24.5331 | QC OK | 4.065 | 0.010145 | 0.12405 | 0.24958 | 3.0516 | 0.72415 | 10.3285 | 12.2269 | 14.2631 | QC OK | N | N
Acquisition-0115.cf | 2023/10/03 10:29:10 | WS23-139-10MC-3 5-7.5cm | WS23-139-10MC-3 5-7.5cm | 0 | 0 | 936 | 4.6265 | 3230 | -25.8169 | 13C QC OK; DO NOT USE 15N | 16.815 | 0.018881 | 0.12502 | 0.11229 | 0.7435 | 1.3477 | 10.4096 | 6.6212 | 7.7239 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0116.cf | 2023/10/03 10:36:17 | WS23-139-18MC-3 0-2.5cm | WS23-139-18MC-3 0-2.5cm | 0 | 0 | 936 | 5.0125 | 3043 | -22.99 | 13C QC OK; DO NOT USE 15N | 17.304 | 0.018746 | 0.11734 | 0.10833 | 0.67809 | 1.338 | 9.7699 | 6.2593 | 7.3017 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0117.cf | 2023/10/03 10:43:24 | WS23-139-01MC-3 17.5-20cm | WS23-139-01MC-3 17.5-20cm | 0 | 0 | 288 | 4.7768 | 603 | -25.6501 | 13C QC OK; DO NOT USE 15N | 16.608 | 0.0047092 | 0.022542 | 0.028355 | 0.13573 | 0.33613 | 1.877 | 4.7869 | 5.5841 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0118.cf | 2023/10/03 10:50:31 | WS23-139-01MC-3 15-17.5cm | WS23-139-01MC-3 15-17.5cm | 0 | 0 | 346 | 4.302 | 801 | -25.4707 | 13C QC OK; DO NOT USE 15N | 18.809 | 0.005988 | 0.03044 | 0.031836 | 0.16184 | 0.42741 | 2.5346 | 5.0835 | 5.93 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0119.cf | 2023/10/03 10:57:39 | WS23-139-16MC-3 12.5-15cm | WS23-139-16MC-3 12.5-15cm | 0 | 0 | 885 | 5.0606 | 2788 | -23.5386 | 13C QC OK; DO NOT USE 15N | 16.565 | 0.017552 | 0.10758 | 0.10596 | 0.64946 | 1.2528 | 8.9578 | 6.1294 | 7.1501 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0120.cf | 2023/10/03 11:04:46 | WS23-139-16MC-3 20-22.5cm | WS23-139-16MC-3 20-22.5cm | 0 | 0 | 943 | 5.1821 | 2991 | -23.5546 | 13C QC OK; DO NOT USE 15N | 17.78 | 0.018861 | 0.11531 | 0.10608 | 0.64855 | 1.3462 | 9.6014 | 6.1139 | 7.132 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0121.cf | 2023/10/03 11:11:54 | WS23-139-16MC-3 17.5-20cm | WS23-139-16MC-3 17.5-20cm | 0 | 0 | 1003 | 5.1417 | 3189 | -23.4979 | 13C QC OK; DO NOT USE 15N | 18.789 | 0.019999 | 0.12264 | 0.10644 | 0.65275 | 1.4275 | 10.2119 | 6.1324 | 7.1536 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0122.cf | 2023/10/03 11:19:01 | WS23-139-18MC-3 20-22.5cm | WS23-139-18MC-3 20-22.5cm | 0 | 0 | 967 | 4.8509 | 3227 | -23.0488 | 13C QC OK; DO NOT USE 15N | 18.639 | 0.019427 | 0.12544 | 0.10423 | 0.67302 | 1.3867 | 10.445 | 6.4572 | 7.5325 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0123.cf | 2023/10/03 11:26:08 | WS23-139-16MC-3 10-12.5cm | WS23-139-16MC-3 10-12.5cm | 0 | 0 | 900 | 5.025 | 2852 | -23.5544 | 13C QC OK; DO NOT USE 15N | 16.701 | 0.017805 | 0.10965 | 0.10661 | 0.65656 | 1.2709 | 9.13 | 6.1585 | 7.1841 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0124.cf | 2023/10/03 11:33:16 | WS23-139-16MC-3 27.5-30cm | WS23-139-16MC-3 27.5-30cm | 0 | 0 | 1004 | 4.9667 | 3171 | -23.4367 | 13C QC OK; DO NOT USE 15N | 18.643 | 0.02009 | 0.12199 | 0.10776 | 0.65436 | 1.434 | 10.1575 | 6.0722 | 7.0834 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0125.cf | 2023/10/03 11:40:23 | WS23-139-17MC-3 45-47.5cm | WS23-139-17MC-3 45-47.5cm | 0 | 0 | 1550 | 4.9085 | 9124 | -26.871 | 13C QC OK; DO NOT USE 15N | 18.472 | 0.031744 | 0.36142 | 0.17185 | 1.9566 | 2.2658 | 30.093 | 11.3854 | 13.2814 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0126.cf | 2023/10/03 11:47:30 | WS23-139-01MC-3 7.5-10cm | WS23-139-01MC-3 7.5-10cm | 0 | 0 | 642 | 4.2587 | 3112 | -26.7986 | 13C QC OK; DO NOT USE 15N | 18.293 | 0.012449 | 0.12008 | 0.068054 | 0.65645 | 0.88859 | 9.9986 | 9.646 | 11.2523 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0127.cf | 2023/10/03 11:54:37 | WS23-139-01MC-3 20-22.5cm | WS23-139-01MC-3 20-22.5cm | 0 | 0 | 267 | 3.8766 | 575 | -25.1163 | 13C QC OK; DO NOT USE 15N | 16.12 | 0.0042932 | 0.021494 | 0.026633 | 0.13334 | 0.30644 | 1.7897 | 5.0065 | 5.8402 | C QC OK; N and C:N QC NOK USE CAUTION N <LOQ | N | N
Acquisition-0128.cf | 2023/10/03 12:01:44 | WS23-139-18MC-3 10-12.5cm | WS23-139-18MC-3 10-12.5cm | 0 | 0 | 824 | 4.7839 | 2611 | -23.0869 | 13C QC OK; DO NOT USE 15N | 17.429 | 0.016437 | 0.10152 | 0.094311 | 0.58248 | 1.1733 | 8.453 | 6.1762 | 7.2047 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0129.cf | 2023/10/03 12:08:52 | WS23-139-17MC-3 47.5-48.75cm | WS23-139-17MC-3 47.5-48.75cm | 0 | 0 | 1294 | 5.0464 | 6054 | -26.1607 | 13C QC OK; DO NOT USE 15N | 16.697 | 0.026758 | 0.24203 | 0.16026 | 1.4496 | 1.9099 | 20.1527 | 9.0453 | 10.5515 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0130.cf | 2023/10/03 12:15:59 | WS23-139-18MC-3 5-7.5cm | WS23-139-18MC-3 5-7.5cm | 0 | 0 | 1028 | 5.0278 | 3550 | -22.821 | 13C QC OK; DO NOT USE 15N | 16.244 | 0.020788 | 0.13736 | 0.12797 | 0.84558 | 1.4838 | 11.4368 | 6.6075 | 7.7079 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0131.cf | 2023/10/03 12:23:05 | WS23-139-05MC-3 25-27.5cm | WS23-139-05MC-3 25-27.5cm | 0 | 0 | 637 | 4.4392 | 2322 | -26.367 | 13C QC OK; DO NOT USE 15N | 17.952 | 0.012453 | 0.089876 | 0.069366 | 0.50064 | 0.88883 | 7.4834 | 7.2174 | 8.4193 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0132.cf | 2023/10/03 12:30:13 | WS23-139-05MC-3 52.5-55cm | WS23-139-05MC-3 52.5-55cm | 0 | 0 | 776 | 4.5607 | 2693 | -26.3466 | 13C QC OK; DO NOT USE 15N | 19.005 | 0.015525 | 0.1048 | 0.08169 | 0.55143 | 1.1081 | 8.726 | 6.7503 | 7.8745 | C QC OK; N and C:N use discretion, these are acidified samples | N | N
Acquisition-0133.cf | 2023/10/03 12:37:20 | NIST 2702 | NIST 2702 | 0 | 0 | 620 | 5.9547 | 3678 | -24.1234 | QC OK | 4.674 | 0.011961 | 0.14396 | 0.2559 | 3.08 | 0.85373 | 11.9866 | 12.036 | 14.0403 | QC OK | N | N
Acquisition-0134.cf | 2023/10/03 12:44:28 | NIST 2702 | NIST 2702 | 0 | 0 | 1617 | 5.9117 | 9748 | -24.2066 | QC OK | 12.879 | 0.033338 | 0.39172 | 0.25885 | 3.0416 | 2.3796 | 32.6164 | 11.7502 | 13.7069 | QC OK | N | N
Acquisition-0135.cf | 2023/10/03 12:51:34 | NIST 2702 | NIST 2702 | 0 | 0 | 937 | 5.9966 | 5821 | -24.1714 | QC OK | 7.286 | 0.018848 | 0.23206 | 0.25869 | 3.185 | 1.3453 | 19.3224 | 12.3122 | 14.3626 | QC OK | N | N
Acquisition-0136.cf | 2023/10/03 12:58:41 | USGS 40 | USGS 40 | 0 | 0 | 2748 | -4.5264 | 6587 | -25.8689 | QC OK | 0.624 | 0.058367 | 0.25893 | 9.3537 | 41.4953 | 4.1661 | 21.5596 | 4.4362 | 5.175 | QC OK | N | N
Acquisition-0137.cf | 2023/10/03 13:05:47 | USGS 41a | USGS 41a | 0 | 0 | 3288 | 47.225 | 7771 | 36.8861 | QC OK | 0.734 | 0.069639 | 0.30439 | 9.4876 | 41.4699 | 4.9707 | 25.3446 | 4.3709 | 5.0988 | QC OK | N | N
Acquisition-0138.cf | 2023/10/03 13:12:54 | BLANK | BLANK | 0 | 0 | 70 | 21.6003 | 40 | -15.4134 | QC OK | 0 | -0.00018117 | 0.00023384 | #DIV/0! | #DIV/0! | -0.012931 | 0.01947 | -1.2907 | -1.5057 | QC OK | N | N